{"text_format":"1","workbook":"xlsx","n_sheets":10,"sheets":["Ca-ion","Ca,ioPassing-Bablok 725-835plab","Ca,ionPassing-Bablok725-835MB","Ca,ionPassing-Bablok 725-837HOV","Ca,ion Bias plot 725-835plab","Ca,ion Bias plot 725-835MB","Ca,ion Bias plot 725-837HOV","Ca,ion Descriptives ","Ark3","Ark2"],"not_rendered":{"charts":22,"images":0,"macros":0,"external_links":0,"defined_names":597,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Ca-ion","index":0,"n_rows":29,"n_cols":7,"rows_kept":29,"cut":false,"hidden":false,"merged":[]}
[]
[]
[null,"Innkjøring Ca-ion",null,"Ref. område:   1,18 - 1,32 mmol/L",null,null,"Måleområdet: 0,20 - 9,99 mmol/L"]
[]
["Dato","Labnummer","735 MB","835 plab","835 MB","837 HOV"]
["2011-10-10",47062399,1.35,1.37,1.35,1.37]
[null,47062351,1.23,1.23,1.22,1.24]
[null,47062015,1.22,1.22,1.21,1.22]
[null,47062418,1.33,1.34,1.32,1.34]
[null,47061998,1.26,1.26,1.25,1.27]
[null,49845,1.22,1.22,1.2,1.22]
[null,47061960,1.1499999999999999,1.1599999999999999,1.1499999999999999,1.1599999999999999]
[null,47061948,1.26,1.25,1.25,1.26]
["2011-10-11",47062097,1.18,1.17,1.1599999999999999,1.18]
[null,47062094,1.25,1.26,1.24,1.26]
[null,47062059,1.25,1.25,1.23,1.25]
[null,47058620,1.34,1.35,1.33,1.35]
[null,49970,1.25,1.24,1.22,1.24]
[null,47062627,1.1200000000000001,1.1200000000000001,1.1100000000000001,1.1299999999999999]
[null,50053,1.1000000000000001,1.0900000000000001,1.08,1.0900000000000001]
["2011-10-14",47063327,1.74,1.78,1.8,1.81]
["2011-10-28",47066402,1.32,1.33,1.33,1.33]
[null,47066959,0.79,0.8,0.81,0.78]
["2011-11-01",47067369,1.1399999999999999,1.1499999999999999,1.1499999999999999,1.1399999999999999]
[null,53820,1.26,1.27,1.27,1.27]
[null,53862,1.17,1.17,1.18,1.19]
[null,53700,1.1599999999999999,1.1399999999999999,1.1499999999999999,1.1499999999999999]
["2011-11-04",47068510,1.1599999999999999,1.1599999999999999,1.18,1.1599999999999999]
[null,47068354,1.17,1.18,1.21,1.17]

{"sheet":"Ca,ioPassing-Bablok 725-835plab","index":1,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-14T08:05:12.150"]
[]
[]
[null,"n ",24," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-5.7500000000000051E-2,-0.11045454545454536,0,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",1.05,1,1.0909090909090908,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","< 0.01 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":"(non-linear relationship between X and Y detected)"}]

{"sheet":"Ca,ionPassing-Bablok725-835MB","index":2,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-14T08:07:30.750"]
[]
[]
[null,"n ",24," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-0.01,-0.12045454545454537,6.0636363636363655E-2,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",1,0.94545454545454544,1.0909090909090908,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"Ca,ionPassing-Bablok 725-837HOV","index":3,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-14T08:09:31.800"]
[]
[]
[null,"n ",24," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-7.1325757575756987E-2,-0.13,4.9999999999998934E-3,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",1.0615530303030298,1,1.1111111111111112,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"Ca,ion Bias plot 725-835plab","index":4,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-14T08:11:53.090"]
[]
[]
[null,null,null,null,"n ",24," "]
[]
[null,null,null,null,"Bias ",3.7499999999999999E-3]
[null,null,null,null,"95% CI ",-1.3556398800936415E-3,8.8556398800936475E-3," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-1.99481977709405E-2,-2.8215338305141061E-2,-1.1681057236739938E-2]
[null,null,null,null,"Upper ",2.7448197770940503E-2,1.9181057236739942E-2,3.5715338305141064E-2]

{"sheet":"Ca,ion Bias plot 725-835MB","index":5,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-14T08:13:30.350"]
[]
[]
[null,null,null,null,"n ",24," "]
[]
[null,null,null,null,"Bias ",-8.3333333333333404E-4]
[null,null,null,null,"95% CI ",-9.5422163738957469E-3,7.8755497072290782E-3," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-4.1256245545358027E-2,-5.5357819820691524E-2,-2.7154671270024525E-2]
[null,null,null,null,"Upper ",3.9589578878691349E-2,2.5488004603357851E-2,5.3691153154024847E-2]

{"sheet":"Ca,ion Bias plot 725-837HOV","index":6,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-14T08:14:35.460"]
[]
[]
[null,null,null,null,"n ",24," "]
[]
[null,null,null,null,"Bias ",6.6666666666666636E-3]
[null,null,null,null,"95% CI ",-1.1549919397474302E-4,1.344883252730807E-2," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-2.481324802677878E-2,-3.5795048469616009E-2,-1.3831447583941548E-2]
[null,null,null,null,"Upper ",3.8146581360112099E-2,2.716478091727487E-2,4.9128381802949324E-2]

{"sheet":"Ca,ion Descriptives ","index":7,"n_rows":35,"n_cols":13,"rows_kept":35,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Comparative descriptives"]
[null,null," "]
[null,"Variables "," 735 MB, 835 plab, 835 MB, 837 HOV"]
[null,"Performed by "," tingstad",null,null,null,null,null,null,null,"Date ","2011-12-14T08:16:42.410"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null," ","n","Mean","SD","SE","95% CI of Mean",null,null,"Median","IQR","95% CI of Median"]
[null,"735 MB ",24,1.2258333333333338,0.15673063757167324,3.1992507425973454E-2,1.1596517960403518,1.2920148706263157,null,1.2250000000000001,0.1,1.1599999999999999,1.26]
[null,"835 plab ",24,1.2295833333333335,0.16334756666761796,3.3343182421743459E-2,1.1606077122127525,1.2985589544539144,null,1.2250000000000001,0.1075,1.1599999999999999,1.26]
[null,"835 MB ",24,1.2250000000000001,0.16333481217904902,3.334057892266646E-2,1.1560297646270556,1.2939702353729445,null,1.2150000000000001,0.1075,1.1599999999999999,1.25]
[null,"837 HOV ",24,1.2324999999999999,0.16968641153866615,3.463709371280399E-2,1.1608477197097284,1.3041522802902719,null,1.23,0.11,1.1599999999999999,1.27]

{"sheet":"Ark3","index":8,"n_rows":0,"n_cols":0,"rows_kept":0,"cut":false,"hidden":false,"merged":[]}

{"sheet":"Ark2","index":9,"n_rows":0,"n_cols":0,"rows_kept":0,"cut":false,"hidden":false,"merged":[]}