{"text_format":"1","workbook":"xlsx","n_sheets":117,"sheets":["pH","pH Passing-Bablok 735-835PLAB","pH Passing-Bablok 735-835MB","pH Passing-Bablok 735-837HOV","pH Bias plot  735-835PLAB","pH Bias plot  735-835MB","pH Bias plot  735-837HOV","pH Correlation 735-835PLAB","pH Correlation 735-835MB","pH Correlation 735-837HOV","pH Descriptives sammeligning","pCO2","pCO2 Passing-Bablok 735-835PLAB","pCO2 Passing-Bablok 735-835MB","pCO2 Passing-Bablok 735-837HOV","pCO2 Bias 735-835PLAB","pCO2 Bias 735-835MB","pCO2 Bias 735-837HOV","pCO2 Correlation 735-835PLAB","pCO2 Correlation 735-835MB","pCO2 Correlation 735-837HOV","pCO2 Descriptives","pO2","pO2 Passing-Bablok 735-835PLAB","pO2 Passing-Bablok 735-835MB","pO2 Passing-Bablok 735-837HOV","pO2 Bias 735-835PLAB","pO2 Bias 735-835MB","pO2 Bias 735-837HOV","pO2 Correlation 735-835PLAB","pO2 Correlation 735-835MB","pO2 Correlation 735-837HOV","pO2 Descriptives","Hb","Hb Passing-Bablok 735-835PLAB","Hb Passing-Bablok 735-835MB","Hb Passing-Bablok 735-837HOV","Hb Bias 735-835PLAB","Hb Bias 735-835MB","Hb Bias 735-837HOV","Hb Descriptives","MetHb","O2Hb","Na","Na Passing-Bablok 735-835plab","Na Passing-Bablok 735-835MB","Na Passing-Bablok 735-835HOV","Na Bias plot 735-835plab","Na Bias plot735-835MB","Na Bias plot 735-835HOV","Na Descriptives ","K","K Passing-Bablok 735-835plab","K Passing-Bablok 735-835MB","K Bias plot 735-835plab","K Bias plot 735-835MB","K Passing-Bablok 735-835HOV","K Bias plot 725-835HOV","K Descriptives  ","Ca","Ca Passing-Bablok725-835plab","Ca Passing-Bablok 725-735MB","Ca Passing-Bablok 725-837HOV","Ca Bias plot 725-835plab","Ca Bias plot 725-835MB","Ca Bias plot 725-837HOV","Ca Descriptives ","Cl","Cl Passing-Bablok725-825MB","Cl Passing-Bablok 725-837HOV","Cl Bias plot 725-835MB","Cl Bias plot 725-837HOV","Cl Descriptives ","Glukose","Glu Passing-Bablok725-835plab","Glu Passing-Bablok 725-835MB","Glu Passing-Bablok 725-837HOV","Glu Bias plot 725-835plab","GluBias plot 725-837HOV","Glu Bias plot 725-835MB","Glu Descriptives ","Lakt","Lakt Passing-Bablok 725-835plab","Lakt Passing-Bablok 725-835MB","Lakt Passing-Bablok 725-837HOV","Lakt Bias plot 725-835plab","Lakt Bias plot 725-835MB","Lakt Bias plot 725-837HOV","Lakt Descriptives ","Kreat","KreatPassing-Bablok Arch-837HOV","Kreat Bias plot Arch-837HOV","Kreat Descriptives ","sO2","sO2Passing-Bablok725-835plab","sO2Passing-Bablok 725-835MB","sO2Passing-Bablok725-837HOV","sO2 Bias plot 725-835plab","sO2 Bias plot 725-835MB","sO2 Bias plot 725-837HOV","sO2 Descriptives ","COHb","COHbPassing-Bablok 725-835plab","COHb Passing-Bablok 725-835MB","COHb Passing-Bablok 725-837HOV","COHb Bias plot 725-835plab","COHbBias plot 725-835MB","COHb Bias plot 725-837HOV","COHb Descriptives ","ctO2","tO2 Passing-Bablok725-835plab","tO2 Passing-Bablok725-835MB","tO2 Passing-Bablok 725-837HOV","tO2 Bias plot 725-835plab","tO2 Bias plot725-835MB","tO2 Bias plot 725-837HOV","tO2 Descriptives "],"not_rendered":{"charts":296,"images":0,"macros":1,"external_links":0,"defined_names":8017,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"pH","index":0,"n_rows":45,"n_cols":6,"rows_kept":45,"cut":false,"hidden":false,"merged":[]}
[]
[]
[null,"Innkjøring pH",null,"Ref. området:  7,35 - 7,45 ",null,"(6,3-8,0)"]
[]
["Dato","Labnummer","735 MB","835 plab","835 MB","837 HOV"]
["2011-10-10",47062477,7.3559999999999999,7.351,7.3529999999999998,7.359]
["2011-10-10",47062478,7.3940000000000001,7.39,7.3879999999999999,7.3890000000000002]
["2011-10-10",47062589,7.4569999999999999,7.4560000000000004,7.4539999999999997,7.4560000000000004]
["2011-10-10",47062611,7.4630000000000001,7.4619999999999997,7.4569999999999999,7.4649999999999999]
["2011-10-10",47062649,7.4610000000000003,7.4660000000000002,7.46,7.4619999999999997]
["2011-10-10",47062666,7.45,7.444,7.4420000000000002,7.4429999999999996]
["2011-10-11",47062193,7.4809999999999999,7.4770000000000003,7.4740000000000002,7.4749999999999996]
["2011-10-11",47062501,7.4139999999999997,7.415,7.4109999999999996,7.4119999999999999]
["2011-10-11",47062700,7.4109999999999996,7.407,7.4050000000000002,7.4029999999999996]
["2011-10-11",47062839,7.4189999999999996,7.4169999999999998,7.4160000000000004,7.4180000000000001]
["2011-10-12",47063235,7.3520000000000003,7.3470000000000004,7.3460000000000001,7.351]
["2011-10-12",47063242,7.4729999999999999,7.4720000000000004,7.468,7.4720000000000004]
["2011-10-13",47063139,7.3970000000000002,7.4,7.3979999999999997,7.3979999999999997]
["2011-10-13",47063273,7.415,7.415,7.4130000000000003,7.41]
["2011-10-13",47063317,7.3689999999999998,7.359,7.367,7.37]
["2011-10-13",47063363,7.4379999999999997,7.4349999999999996,7.431,7.4329999999999998]
["2011-10-14",47063399,7.4,7.3959999999999999,7.3929999999999998,7.3959999999999999]
["2011-10-14",47063445,7.2770000000000001,7.274,7.2709999999999999,7.2720000000000002]
["2011-10-14",47063456,7.4420000000000002,7.4409999999999998,7.4349999999999996,7.4379999999999997]
["2011-10-14",47063598,7.1920000000000002,7.1870000000000003,7.1840000000000002,7.1879999999999997]
["2011-10-14",47063611,7.3550000000000004,7.3460000000000001,7.34,7.3520000000000003]
["2011-10-15",47063791,7.3490000000000002,7.3449999999999998,7.3390000000000004,7.3449999999999998]
["2011-10-16",47063944,7.4020000000000001,7.41,7.4020000000000001,7.3970000000000002]
["2011-10-16",47063949,7.4109999999999996,7.4119999999999999,7.4039999999999999,7.4059999999999997]
["2011-10-16",47063979,7.44,7.4420000000000002,7.4320000000000004,7.4340000000000002]
["2011-10-19",47064495,7.3159999999999998,7.3179999999999996,7.319,7.3280000000000003]
["2011-10-19",47064940,7.4749999999999996,7.468,7.4669999999999996,7.4720000000000004]
["2011-10-20",47065073,7.4829999999999997,7.5030000000000001,7.4770000000000003,7.484]
["2011-10-20",47065180,7.4909999999999997,7.4880000000000004,7.4809999999999999,7.4829999999999997]
["2011-10-21",47065421,7.3609999999999998,7.3570000000000002,7.3520000000000003,7.3550000000000004]
["2011-10-21",47065432,7.4640000000000004,7.4630000000000001,7.46,7.4660000000000002]
["2011-10-26",47066431,7.3979999999999997,7.3979999999999997,7.3979999999999997,7.4059999999999997]
["2011-10-26",47066582,7.4109999999999996,7.4080000000000004,7.4039999999999999,7.4050000000000002]
["2011-10-27",47066702,7.327,7.3209999999999997,7.319,7.3289999999999997]
["2011-10-27",47066712,7.4779999999999998,7.4720000000000004,7.4690000000000003,7.4720000000000004]
["2011-10-28",47066955,7.4889999999999999,7.48,7.476,7.48]
["2011-11-01",47067763,7.4740000000000002,7.4610000000000003,7.4560000000000004,7.4660000000000002]
["2011-11-01",47067808,7.4269999999999996,7.4269999999999996,7.4160000000000004,7.4240000000000004]
["2011-11-04",47068634,7.4640000000000004,7.4640000000000004,7.4619999999999997,7.4610000000000003]
["2011-11-18",47064503,7.3929999999999998,7.3949999999999996,7.3890000000000002,7.39]

{"sheet":"pH Passing-Bablok 735-835PLAB","index":1,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T12:40:02.430"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-5.2123745819413939E-2,-0.25647945205483519,0.12557758620682594,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",1.0066889632107046,0.98275862068966469,1.0342465753424712,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","0.01 < p < 0.05 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":"(non-linear relationship between X and Y detected)"}]

{"sheet":"pH Passing-Bablok 735-835MB","index":2,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T12:41:12.780"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",8.4240909771050393E-2,-7.0717391304311789E-2,0.2847647058823245,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.98776746969923823,0.960784313725494,1.0086956521739081,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"pH Passing-Bablok 735-837HOV","index":3,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-11-29T11:49:35.640"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",0.1085984848484475,-3.5000000000002807E-3,0.30871052631575602,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.98484848484848986,0.9578947368421098,1,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"pH Bias plot  735-835PLAB","index":4,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," fringvaldsen",null,null,"Date ","2011-12-06T13:16:49.520"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-1.9999999999999129E-3]
[null,null,null,null,"95% CI ",-3.7411897371454439E-3,-2.5881026285438203E-4," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-1.26707440885412E-2,-1.55541801826715E-2,-9.7873079944108994E-3]
[null,null,null,null,"Upper ",8.6707440885413715E-3,5.7873079944110719E-3,1.155418018267167E-2]

{"sheet":"pH Bias plot  735-835MB","index":5,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," fringvaldsen",null,null,"Date ","2011-12-06T13:18:21.120"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-6.0249999999999471E-3]
[null,null,null,null,"95% CI ",-7.3750559672085148E-3,-4.6749440327913793E-3," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-1.429871160302592E-2,-1.6534424358447364E-2,-1.2062998847604473E-2]
[null,null,null,null,"Upper ",2.2487116030260228E-3,1.2998847604577361E-5,4.4844243584474686E-3]

{"sheet":"pH Bias plot  735-837HOV","index":6,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," fringvaldsen",null,null,"Date ","2011-11-29T12:09:08.390"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-2.5999999999999357E-3]
[null,null,null,null,"95% CI ",-3.98517303962627E-3,-1.2148269603736014E-3," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-1.1088923813915153E-2,-1.3382790963409455E-2,-8.7950566644208507E-3]
[null,null,null,null,"Upper ",5.8889238139152797E-3,3.5950566644209776E-3,8.1827909634095818E-3]

{"sheet":"pH Correlation 735-835PLAB","index":7,"n_rows":12,"n_cols":12,"rows_kept":12,"cut":false,"hidden":false,"merged":[]}
[null,null,null,null,null,null,null,null,null,null,null,"Beregnet"]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Pearson correlation"]
[null,null," "]
[null,"Alternative hypothesis "," 735 MB  ¹  835 plab"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T12:25:32.550"]
[]
[]
[null,"n ",40," "]
[]
[null,"r statistic ",0.9965194876333846]
[null,"95% CI ",0.99338037901301801,null," "]

{"sheet":"pH Correlation 735-835MB","index":8,"n_rows":14,"n_cols":9,"rows_kept":14,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Pearson correlation"]
[null,null," "]
[null,"Alternative hypothesis "," 735 MB  ¹  835 plab"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T12:23:32.510"]
[]
[]
[null,"n ",40," "]
[]
[null,"r statistic ",0.9965194876333846]
[null,"95% CI ",0.99338037901301801,0.99817135653092648," "]
[]
[null,"2-tailed p ",0," (t approximation)"]

{"sheet":"pH Correlation 735-837HOV","index":9,"n_rows":14,"n_cols":9,"rows_kept":14,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Pearson correlation"]
[null,null," "]
[null,"Alternative hypothesis "," 735 MB  ¹  837 HOV"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-11-29T12:13:41.300"]
[]
[]
[null,"n ",40," "]
[]
[null,"r statistic ",0.99774262959400439]
[null,"95% CI ",0.99570431128318071,0.99881433383196283," "]
[]
[null,"2-tailed p ",0," (t approximation)"]

{"sheet":"pH Descriptives sammeligning","index":10,"n_rows":35,"n_cols":13,"rows_kept":35,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Comparative descriptives"]
[null,null," "]
[null,"Variables "," 735 MB, 835 plab, 835 MB, 837 HOV"]
[null,"Performed by "," fringvaldsen",null,null,null,null,null,null,null,"Date ","2011-11-29T12:11:45.580"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null," ","n","Mean","SD","SE","95% CI of Mean",null,null,"Median","IQR","95% CI of Median"]
[null,"735 MB ",40,7.4117249999999988,6.3489949739692755E-2,1.0038642485352182E-2,7.391419931055915,7.4320300689440826,null,7.4145000000000003,7.675000000000054E-2,7.3979999999999997,7.45]
[null,"835 plab ",40,7.4097249999999972,6.4545745058277953E-2,1.0205578382835796E-2,7.3890822713661102,7.4303677286338843,null,7.415,8.0499999999999794E-2,7.3979999999999997,7.444]
[null,"835 MB ",40,7.4057000000000004,6.3058458835981962E-2,9.9704177830836557E-3,7.3855329285277138,7.425867071472287,null,7.4119999999999999,7.6500000000000234E-2,7.3979999999999997,7.4420000000000002]
[null,"837 HOV ",40,7.4091250000000013,6.2584583790300408E-2,9.8954915595501538E-3,7.3891094811043487,7.4291405188956539,null,7.4109999999999996,8.0000000000000071E-2,7.3970000000000002,7.4429999999999996]

{"sheet":"pCO2","index":11,"n_rows":44,"n_cols":6,"rows_kept":44,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Innkjøring pCO2",null,"Ref. området:  4,50 - 6,0 KPA       (0,67-33,3)"]
[]
["Dato","Labnummer","735 MB","835 plab","835 MB","837 HOV"]
["2011-10-10",47062477,9.3800000000000008,9.27,9.0500000000000007,8.8000000000000007]
["2011-10-10",47062478,5.64,5.64,5.55,5.55]
["2011-10-10",47062589,5.22,5.05,4.97,4.93]
["2011-10-10",47062611,5.03,4.99,4.8499999999999996,4.8099999999999996]
["2011-10-10",47062649,4.6399999999999997,4.5199999999999996,4.58,4.46]
["2011-10-10",47062666,5.53,5.33,5.27,5.31]
["2011-10-11",47062501,5.26,5.14,5.32,5.15]
["2011-10-11",47062700,4.38,4.34,4.49,4.32]
["2011-10-11",47062193,4.0999999999999996,4.0599999999999996,3.91,4.0199999999999996]
["2011-10-11",47062839,6.56,6.24,6.23,6.06]
["2011-10-12",47063235,5.15,5.24,5.12,5.08]
["2011-10-12",47063242,4.9000000000000004,4.74,4.68,4.68]
["2011-10-13",47063363,9.51,9.1999999999999993,9.4499999999999993,9.08]
["2011-10-13",47063317,5.81,5.73,5.69,5.59]
["2011-10-13",47063273,4.5,4.68,4.5999999999999996,4.6399999999999997]
["2011-10-13",47063139,5.86,5.62,5.49,5.67]
["2011-10-14",47063445,7.72,7.86,7.89,7.73]
["2011-10-14",47063456,4.6900000000000004,4.7,4.74,4.7699999999999996]
["2011-10-14",47063399,5.91,6,6.03,5.89]
["2011-10-14",47063598,2.74,2.5299999999999998,2.52,2.46]
["2011-10-14",47063611,5.91,5.78,8.84,5.68]
["2011-10-15",47063791,2.89,2.85,2.79,2.78]
["2011-10-16",47063944,5.0599999999999996,4.76,4.82,5]
["2011-10-16",47063949,5.34,5.17,5.24,5.29]
["2011-10-16",47063979,5,4.59,4.8499999999999996,4.93]
["2011-11-18",47064503,4.6100000000000003,4.34,4.53,4.4400000000000004]
["2011-10-19",47064495,7.22,7.09,6.97,6.79]
["2011-10-19",47064940,4.66,4.66,4.58,4.54]
["2011-10-20",47065180,4.0599999999999996,3.93,4.01,4.03]
["2011-10-20",47065073,5.74,5.31,5.81,5.66]
["2011-10-21",47065432,4.9000000000000004,4.82,4.67,4.84]
["2011-10-21",47065421,6.74,6.55,6.62,6.54]
["2011-10-26",47066431,5.99,5.86,5.9,5.83]
["2011-10-26",47066582,5.81,5.34,5.48,5.61]
["2011-10-27",47066702,6.13,6.02,6.04,5.78]
["2011-10-27",47066712,4.41,4.0999999999999996,4.08,4.08]
["2011-10-28",47066955,3.86,3.83,3.89,3.92]
["2011-11-01",47067763,4.17,4.17,4.17,4.1100000000000003]
["2011-11-01",47067808,5.61,5.3,5.5,5.52]
["2011-11-04",47068634,5.49,5.19,5.0999999999999996,5.31]

{"sheet":"pCO2 Passing-Bablok 735-835PLAB","index":12,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T13:26:45.040"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-1.5725937619592921E-2,-0.23625000000000085,0.20151639344262362,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.98043436662839623,0.94057377049180313,1.017857142857143,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"pCO2 Passing-Bablok 735-835MB","index":13,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T13:29:03.580"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-5.7175198352152623E-2,-0.32557692307692743,0.17483050847457626,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.99126487641135208,0.94703389830508478,1.0384615384615394,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"pCO2 Passing-Bablok 735-837HOV","index":14,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-11-29T13:11:07.080"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",0.17444216887695219,-6.3960000000001641E-2,0.37699999999999945,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.93786301829780072,0.9,0.98400000000000032,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":"(proportional bias detected)"}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"pCO2 Bias 735-835PLAB","index":15,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," fringvaldsen",null,null,"Date ","2011-12-06T13:59:19.510"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-0.13975000000000004]
[null,null,null,null,"95% CI ",-0.1885085168886006,-9.0991483111399485E-2," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.43856272830612175,-0.51930755725470967,-0.35781789935753383]
[null,null,null,null,"Upper ",0.15906272830612159,7.8317899357533666E-2,0.23980755725470951]

{"sheet":"pCO2 Bias 735-835MB","index":16,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," fringvaldsen",null,null,"Date ","2011-12-06T14:00:19.450"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-4.5250000000000012E-2]
[null,null,null,null,"95% CI ",-0.20641888915580561,0.1159188891558056," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-1.0329608361754239,-1.2998589130068368,-0.76606275934401102]
[null,null,null,null,"Upper ",0.94246083617542353,0.67556275934401078,1.2093589130068363]

{"sheet":"pCO2 Bias 735-837HOV","index":17,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," fringvaldsen",null,null,"Date ","2011-11-29T13:15:04.370"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-0.16125]
[null,null,null,null,"95% CI ",-0.2102965069374203,-0.11220349306257982," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.4618276526250758,-0.54304939737657809,-0.38060590787357351]
[null,null,null,null,"Upper ",0.13932765262507563,5.8105907873573344E-2,0.22054939737657792]

{"sheet":"pCO2 Correlation 735-835PLAB","index":18,"n_rows":14,"n_cols":9,"rows_kept":14,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Pearson correlation"]
[null,null," "]
[null,"Alternative hypothesis "," 735 MB  ¹  835 plab"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T14:03:51.380"]
[]
[]
[null,"n ",40," "]
[]
[null,"r statistic ",0.99383031702349711]
[null,"95% CI ",0.98828005625301463,0.99675640649206132," "]
[]
[null,"2-tailed p ",0," (t approximation)"]

{"sheet":"pCO2 Correlation 735-835MB","index":19,"n_rows":14,"n_cols":9,"rows_kept":14,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Pearson correlation"]
[null,null," "]
[null,"Alternative hypothesis "," 735 MB  ¹  835 MB"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T14:06:52.570"]
[]
[]
[null,"n ",40," "]
[]
[null,"r statistic ",0.93988677457376313]
[null,"95% CI ",0.88852203728113033,0.96798582625107654," "]
[]
[null,"2-tailed p ",0," (t approximation)"]

{"sheet":"pCO2 Correlation 735-837HOV","index":20,"n_rows":14,"n_cols":9,"rows_kept":14,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Pearson correlation"]
[null,null," "]
[null,"Alternative hypothesis "," 735 MB  ¹  837 HOV"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T14:07:41.750"]
[]
[]
[null,"n ",40," "]
[]
[null,"r statistic ",0.99504356467868171]
[null,"95% CI ",0.99057958742784047,0.99739500016189209," "]
[]
[null,"2-tailed p ",0," (t approximation)"]

{"sheet":"pCO2 Descriptives","index":21,"n_rows":35,"n_cols":13,"rows_kept":35,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Comparative descriptives"]
[null,null," "]
[null,"Variables "," 735 MB, 835 plab, 835 MB, 837 HOV"]
[null,"Performed by "," fringvaldsen",null,null,null,null,null,null,null,"Date ","2011-11-29T13:23:35.510"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null," ","n","Mean","SD","SE","95% CI of Mean",null,null,"Median","IQR","95% CI of Median"]
[null,"735 MB ",40,5.4032499999999999,1.3739301432247342,0.21723742992757594,4.9638458675835606,5.8426541324164392,null,5.24,1.2649999999999999,4.9000000000000004,5.74]
[null,"835 plab ",40,5.2635000000000005,1.3607040975265861,0.21514620848539495,4.8283257617772026,5.6986742382227984,null,5.1550000000000002,1.1950000000000001,4.7,5.34]
[null,"835 MB ",40,5.3579999999999988,1.475269988328082,0.23326066634033796,4.8861858160707961,5.8298141839292015,null,5.1100000000000003,1.2975000000000001,4.68,5.5]
[null,"837 HOV ",40,5.2420000000000009,1.2974574742302338,0.20514603928883796,4.8270530111572736,5.6569469888427282,null,5.1150000000000002,1.1575,4.7699999999999996,5.59]

{"sheet":"pO2","index":22,"n_rows":44,"n_cols":6,"rows_kept":44,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Innkjøring pO2",null,"Ref. området:  11 - 14 KPA        (0,00-107)"]
[]
["Dato","Labnummer","735 MB","835 plab","835 MB","837 HOV"]
["2011-10-10",47062477,7.55,7.58,7.55,7.82]
["2011-10-10",47062478,12.9,13.5,13.5,13.9]
["2011-10-10",47062589,10.7,11.5,11.5,12.2]
["2011-10-10",47062611,8.56,9.65,9.51,9.8699999999999992]
["2011-10-10",47062649,11.6,13.6,12.3,14.8]
["2011-10-10",47062666,8.73,8.8000000000000007,8.6,8.86]
["2011-10-11",47062501,11.5,12,11.8,12]
["2011-10-11",47062700,11.8,11.9,11.7,12]
["2011-10-11",47062193,8.69,8.91,8.5,8.76]
["2011-10-11",47062839,6.19,6.56,6.22,6.38]
["2011-10-12",47063235,13.1,13.8,13.6,14.3]
["2011-10-12",47063242,9.6,9.94,9.7799999999999994,10.3]
["2011-10-13",47063363,5.9,6.06,5.84,6.04]
["2011-10-13",47063317,7.45,7.42,7.67,7.78]
["2011-10-13",47063273,10.199999999999999,10.8,10.6,11.2]
["2011-10-13",47063139,11.1,11.7,11.8,12.4]
["2011-10-14",47063445,9.41,9.43,9.31,9.58]
["2011-10-14",47063456,9.91,9.32,9.1199999999999992,9.42]
["2011-10-14",47063399,10.199999999999999,10.199999999999999,9.9700000000000006,10.3]
["2011-10-14",47063598,16.3,17.399999999999999,17.600000000000001,18.3]
["2011-10-14",47063611,9.9499999999999993,10,10.1,10.4]
["2011-10-15",47063791,11.4,12.4,12.2,13.2]
["2011-10-16",47063944,11.3,12.7,12,11.8]
["2011-10-16",47063949,9.9700000000000006,10.1,9.85,10.1]
["2011-10-16",47063979,6.49,6.82,6.41,6.61]
["2011-11-18",47064503,5.85,5.96,5.81,5.99]
["2011-10-19",47064495,16.100000000000001,16.600000000000001,16.7,17.5]
["2011-10-19",47064940,8.08,8.09,8.08,8.33]
["2011-10-20",47065180,11.7,12.1,12,12.1]
["2011-10-20",47065073,7.64,8.5500000000000007,7.92,8.0500000000000007]
["2011-10-21",47065432,15.7,16.399999999999999,16.600000000000001,17.399999999999999]
["2011-10-21",47065421,11.2,11.3,11.1,11.4]
["2011-10-26",47066431,12.1,12.7,12.9,13.3]
["2011-10-26",47066582,11.2,11.9,11.2,11.2]
["2011-10-27",47066702,7.71,7.99,7.9,8.34]
["2011-10-27",47066712,14.2,14.5,14.5,15.3]
["2011-10-28",47066955,16.8,17.7,17,17.3]
["2011-11-01",47067763,8.48,8.56,8.51,8.82]
["2011-11-01",47067808,10.3,10.6,10.1,10.3]
["2011-11-04",47068634,11.6,12.7,12.5,11.8]

{"sheet":"pO2 Passing-Bablok 735-835PLAB","index":23,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T14:11:52.270"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-0.45024924723987458,-1.0144505494505409,-4.764462809916846E-2,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":"(constant bias detected)"}]
[null,"Slope ",1.0817999330879888,1.0413223140495862,1.131868131868131,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":"(proportional bias detected)"}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"pO2 Passing-Bablok 735-835MB","index":24,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T14:12:39.020"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-0.63957142857142979,-1.1285641547861536,-0.21311355311355246,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":"(constant bias detected)"}]
[null,"Slope ",1.0857142857142859,1.0402930402930401,1.1364562118126276,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":"(proportional bias detected)"}]
[]
[null,"Cusum test for linearity - p ","0.05 < p < 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":"(non-linear relationship between X and Y detected)"}]

{"sheet":"pO2 Passing-Bablok 735-837HOV","index":25,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-11-29T14:54:51.680"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-0.71648063781321292,-1.3440298507462667,-0.10967940813810106,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":"(constant bias detected)"}]
[null,"Slope ",1.1298405466970389,1.0530209617755857,1.1940298507462686,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":"(proportional bias detected)"}]
[]
[null,"Cusum test for linearity - p ","0.05 < p < 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":"(non-linear relationship between X and Y detected)"}]

{"sheet":"pO2 Bias 735-835PLAB","index":26,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," fringvaldsen",null,null,"Date ","2011-12-06T14:17:01.600"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",0.46449999999999997]
[null,null,null,null,"95% CI ",0.31178341637997925,0.61721658362002063," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.47141154778863326,-0.72431248097707135,-0.21851061460019519]
[null,null,null,null,"Upper ",1.400411547788633,1.147510614600195,1.653312480977071]

{"sheet":"pO2 Bias 735-835MB","index":27,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," fringvaldsen",null,null,"Date ","2011-12-06T14:47:00.650"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",0.26724999999999999]
[null,null,null,null,"95% CI ",0.12953783902846705,0.40496216097153293," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.57670812602036348,-0.80476151220501413,-0.34865473983571282]
[null,null,null,null,"Upper ",1.1112081260203632,0.88315473983571258,1.3392615122050138]

{"sheet":"pO2 Bias 735-837HOV","index":28,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," fringvaldsen",null,null,"Date ","2011-11-29T13:41:27.290"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",0.65725000000000011]
[null,null,null,null,"95% CI ",0.42907550401311179,0.88542449598688844," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.74109941722061934,-1.1189597752578624,-0.36323905918337634]
[null,null,null,null,"Upper ",2.0555994172206193,1.6777390591833765,2.4334597752578624]

{"sheet":"pO2 Correlation 735-835PLAB","index":29,"n_rows":14,"n_cols":9,"rows_kept":14,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Pearson correlation"]
[null,null," "]
[null,"Alternative hypothesis "," 735 MB  ¹  835 plab"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T14:48:50.440"]
[]
[]
[null,"n ",40," "]
[]
[null,"r statistic ",0.98950879623055388]
[null,"95% CI ",0.98010966853488235,0.99447877354913528," "]
[]
[null,"2-tailed p ",0," (t approximation)"]

{"sheet":"pO2 Correlation 735-835MB","index":30,"n_rows":14,"n_cols":9,"rows_kept":14,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Pearson correlation"]
[null,null," "]
[null,"Alternative hypothesis "," 735 MB  ¹  835 MB"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T14:49:47.800"]
[]
[]
[null,"n ",40," "]
[]
[null,"r statistic ",0.99254566203197048]
[null,"95% CI ",0.98584792565238033,0.99607982563632269," "]
[]
[null,"2-tailed p ",0," (t approximation)"]

{"sheet":"pO2 Correlation 735-837HOV","index":31,"n_rows":14,"n_cols":9,"rows_kept":14,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Pearson correlation"]
[null,null," "]
[null,"Alternative hypothesis "," 735 MB  ¹  837 HOV"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T14:50:28.300"]
[]
[]
[null,"n ",40," "]
[]
[null,"r statistic ",0.98137355449823871]
[null,"95% CI ",0.96481491634734839,0.99017839837885857," "]
[]
[null,"2-tailed p ",0," (t approximation)"]

{"sheet":"pO2 Descriptives","index":32,"n_rows":35,"n_cols":13,"rows_kept":35,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Comparative descriptives"]
[null,null," "]
[null,"Variables "," 735 MB, 835 plab, 835 MB, 837 HOV"]
[null,"Performed by "," fringvaldsen",null,null,null,null,null,null,null,"Date ","2011-11-29T13:40:14.980"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null," ","n","Mean","SD","SE","95% CI of Mean",null,null,"Median","IQR","95% CI of Median"]
[null,"735 MB ",40,10.479000000000001,2.7943567674135212,0.44182659900660526,9.5853214406480181,11.372678559351984,null,10.25,3.1349999999999998,9.41,11.4]
[null,"835 plab ",40,10.943499999999998,3.0200700874869333,0.47751500849013462,9.9776348261605357,11.909365173839461,null,10.7,3.96,9.43,12]
[null,"835 MB ",40,10.74625,3.0370335454864925,0.48019716670369511,9.7749596496464068,11.717540350353593,null,10.35,3.7675000000000001,9.31,11.8]
[null,"837 HOV ",40,11.13625,3.2122471465890094,0.50790086952990221,10.108923627174908,12.163576372825093,null,10.8,4.1950000000000003,9.58,12]

{"sheet":"Hb","index":33,"n_rows":44,"n_cols":6,"rows_kept":44,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Innkjøring Hb",null,"Ref.området: 11,7 - 17,0          (0,00-27,7)"]
[]
["Dato","Labnummer","735 MB","835 plab","835 MB","837 HOV"]
["2011-10-10",47062477,13.1,12.8,12.8,12.6]
["2011-10-10",47062478,15.3,14.9,15.2,14.8]
["2011-10-10",47062589,13.8,13.4,13.5,13.2]
["2011-10-10",47062611,9.1999999999999993,8.9,8.9,8.9]
["2011-10-10",47062649,13.8,13.2,13.7,13.2]
["2011-10-10",47062666,15.5,15.2,15.2,15.1]
["2011-10-11",47062501,9,8.8000000000000007,8.8000000000000007,8.9]
["2011-10-11",47062700,10.3,10.1,10.1,10.1]
["2011-10-11",47062193,16.7,16.399999999999999,16.399999999999999,16.2]
["2011-10-11",47062839,13.2,13.1,13.3,12.9]
["2011-10-12",47063235,9.1,8.8000000000000007,8.9,8.6]
["2011-10-12",47063242,13.3,13,13,12.7]
["2011-10-13",47063363,12.9,12.8,12.7,12.5]
["2011-10-13",47063317,10.7,10.6,10.3,10.4]
["2011-10-13",47063273,10,9.9,10.1,9.9]
["2011-10-13",47063139,9.5,9.4,9.4,9.1999999999999993]
["2011-10-14",47063445,12,11.9,12,11.9]
["2011-10-14",47063456,8.4,8.3000000000000007,8.3000000000000007,8.4]
["2011-10-14",47063399,9.8000000000000007,9.8000000000000007,10,9.6999999999999993]
["2011-10-14",47063598,14.6,14.2,14.3,14.1]
["2011-10-14",47063611,12.2,11.4,12.4,10.8]
["2011-10-15",47063791,4.7,4.7,4.5,4.4000000000000004]
["2011-10-16",47063944,12.4,11.9,12.1,11.9]
["2011-10-16",47063949,13.3,13,13.1,13.2]
["2011-10-16",47063979,13.7,13.4,13.4,13.1]
["2011-11-18",47064503,10.8,10.199999999999999,10.5,10.3]
["2011-10-19",47064495,8.4,8.5,8.4,8.4]
["2011-10-19",47064940,14.2,13.9,14.2,13.8]
["2011-10-20",47065180,11.7,11.4,11.5,11.4]
["2011-10-20",47065073,6.9,7.4,6.1,7.6]
["2011-10-21",47065432,9.5,9.6999999999999993,9.9,9.9]
["2011-10-21",47065421,14.5,14.2,14.3,14.2]
["2011-10-26",47066431,14.3,14,14.2,14]
["2011-10-26",47066582,14.1,13.6,13.5,13.6]
["2011-10-27",47066702,11.6,11.3,11.5,11.2]
["2011-10-27",47066712,15.1,14.7,15,14.5]
["2011-10-28",47066955,15.2,14.8,15.1,15]
["2011-11-01",47067763,11.2,10.9,11.2,11]
["2011-11-01",47067808,12.6,12.3,12.5,12.3]
["2011-11-04",47068634,14.4,14,14.2,14.4]

{"sheet":"Hb Passing-Bablok 735-835PLAB","index":34,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T14:51:31.140"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",0.21799999999999783,-8.9062499999999822E-2,0.52187500000000453,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.96,0.9375,0.984375,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":"(proportional bias detected)"}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"Hb Passing-Bablok 735-835MB","index":35,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T14:53:25.270"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-0.1999999999999944,-0.20000000000000531,0.1969696969697014,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",1,0.96969696969696928,1,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"Hb Passing-Bablok 735-837HOV","index":36,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," fringvaldsen",null,null,null,"Date ","2011-12-06T14:54:34.930"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",0.21304347826087902,-0.29999999999998739,0.56341463414634374,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.95652173913043415,0.92682926829268275,0.99999999999999867,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":"(proportional bias detected)"}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"Hb Bias 735-835PLAB","index":37,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," fringvaldsen",null,null,"Date ","2011-12-06T14:56:18.920"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-0.255]
[null,null,null,null,"95% CI ",-0.32776714588730749,-0.18223285411269233," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.70094771910955134,-0.82145119465971239,-0.58044424355939028]
[null,null,null,null,"Upper ",0.19094771910955141,7.0444243559390343E-2,0.31145119465971249]

{"sheet":"Hb Bias 735-835MB","index":38,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," fringvaldsen",null,null,"Date ","2011-12-06T14:58:02"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-0.16250000000000001]
[null,null,null,null,"95% CI ",-0.22954106666118804,-9.545893333881203E-2," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.57335589381956154,-0.68437689703129134,-0.46233489060783167]
[null,null,null,null,"Upper ",0.24835589381956139,0.13733489060783155,0.35937689703129122]

{"sheet":"Hb Bias 735-837HOV","index":39,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," fringvaldsen",null,null,"Date ","2011-12-06T14:58:51.240"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-0.3175]
[null,null,null,null,"95% CI ",-0.41989411521453929,-0.21510588478546075," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.94501426585949255,-1.1145804431332917,-0.77544808858569325]
[null,null,null,null,"Upper ",0.31001426585949238,0.14044808858569313,0.47958044313329162]

{"sheet":"Hb Descriptives","index":40,"n_rows":35,"n_cols":13,"rows_kept":35,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Comparative descriptives"]
[null,null," "]
[null,"Variables "," 735 MB, 835 plab, 835 MB, 837 HOV"]
[null,"Performed by "," fringvaldsen",null,null,null,null,null,null,null,"Date ","2011-12-06T14:59:55.790"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null," ","n","Mean","SD","SE","95% CI of Mean",null,null,"Median","IQR","95% CI of Median"]
[null,"735 MB ",40,12.025,2.6498911443534956,0.41898457838335407,11.177523783669404,12.872476216330597,null,12.5,4.2249999999999996,10.8,13.7]
[null,"835 plab ",40,11.77,2.5209074475224909,0.39859046523262315,10.96377476697549,12.576225233024509,null,12.1,3.95,10.6,13.2]
[null,"835 MB ",40,11.862500000000001,2.6478522985377775,0.41866220855457542,11.015675838128796,12.709324161871201,null,12.45,4,10.5,13.4]
[null,"837 HOV ",40,11.7075,2.5194003661840259,0.39835217475018903,10.901756754921728,12.513243245078272,null,12.1,3.85,10.4,13.2]

{"sheet":"MetHb","index":41,"n_rows":44,"n_cols":8,"rows_kept":44,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Innkjøring MetHb",null,"Ref. området:  0,0 - 1,5     (0,0-100)"]
[]
["Dato","Labnummer","735 MB","835 plab","835 MB","837 HOV"]
["2011-10-10",47062477,1.1000000000000001,0.2,0.4,0.2,null,1]
["2011-10-10",47062478,1.4,0.5,0.8,0.7,null,2]
["2011-10-10",47062589,1.4,0.7,0.9,0.7,null,3]
["2011-10-10",47062611,1.3,0.7,0.5,0.5,null,4]
["2011-10-10",47062649,1.1000000000000001,0.3,0.6,0.7,null,5]
["2011-10-10",47062666,1.3,0.3,0.6,0.6,null,6]
["2011-10-11",47062501,1.5,0.8,1,0.8,null,7]
["2011-10-11",47062700,1.2,0.7,0.9,1,null,8]
["2011-10-11",47062193,1.3,0.5,0.6,0.5,null,9]
["2011-10-11",47062839,1.1000000000000001,0.4,0.5,0.2,null,10]
["2011-10-12",47063235,1.4,0.5,0.4,0.2," ",11]
["2011-10-12",47063242,1.4,0.5,0.5,0.7,null,12]
["2011-10-13",47063363,1,0.3,0.3,0.3]
["2011-10-13",47063317,1.1000000000000001,0.7,0.6,0.3]
["2011-10-13",47063273,0.9,0.7,0.9,0.5]
["2011-10-13",47063139,0.9,0.6,0.9,0.4]
["2011-10-14",47063445,1.4,0.7,0.9,1]
["2011-10-14",47063456,1,0.8,0.7,0.9]
["2011-10-14",47063399,1,0.8,0.7,1.1000000000000001]
["2011-10-14",47063598,1.5,0.5,0.6,0.6]
["2011-10-14",47063611,1.1000000000000001,0.9,0.6,1]
["2011-10-15",47063791,1.5,1.5,0.4,0.8]
["2011-10-16",47063944,1.2,0.3,0.7,0.3]
["2011-10-16",47063949,1,0.1,0.5,0.3]
["2011-10-16",47063979,1,0.4,0.4,0.5]
["2011-11-18",47064503,1.1000000000000001,0.7,0.7,0.7]
["2011-10-19",47064495,1.1000000000000001,0.7,0.7,0.2]
["2011-10-19",47064940,1.3,0.5,0.6,1]
["2011-10-20",47065180,0.9,0.4,0.6,0.7]
["2011-10-20",47065073,1,0.1,1.3,0.6]
["2011-10-21",47065432,1.1000000000000001,0,0.9,0.2]
["2011-10-21",47065421,1.3,0.4,0.5,0.5]
["2011-10-26",47066431,1.3,0.2,0.7,0.6]
["2011-10-26",47066582,1.5,0.8,1.1000000000000001,0.8]
["2011-10-27",47066702,0.8,0.6,0.4,0.4]
["2011-10-27",47066712,1.2,0.7,0.8,0.8]
["2011-10-28",47066955,1,0.3,0.6,0.6]
["2011-11-01",47067763,0.9,0.2,0.6,0.4]
["2011-11-01",47067808,1.1000000000000001,0.2,0.5,0.3]
["2011-11-04",47068634,1.1000000000000001,0.6,0.5,0.8]

{"sheet":"O2Hb","index":42,"n_rows":44,"n_cols":8,"rows_kept":44,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Innkjøring O2Hb",null,"Ref. området:  94 - 98 %        (0,0-100)"]
[]
["Dato","Labnummer","735 MB","835 plab","835 MB","837 HOV"]
["2011-10-10",47062477,86.1,85.4,85.5,86.4,null,1]
["2011-10-10",47062478,93.1,94.5,94.2,94.5,null,2]
["2011-10-10",47062589,94.8,95.6,95.6,96,null,3]
["2011-10-10",47062611,91,92.5,92.8,93.3,null,4]
["2011-10-10",47062649,95.5,96.9,96.3,96.8,null,5]
["2011-10-10",47062666,91,91.1,90.6,91.2,null,6]
["2011-10-11",47062501,94.6,95.4,95.1,95.5,null,7]
["2011-10-11",47062700,94.4,94.9,94.4,94.6,null,8]
["2011-10-11",47062193,91.7,92.3,92,91.1,null,9]
["2011-10-11",47062839,81,80.599999999999994,80.3,82,null,10]
["2011-10-12",47063235,89.2,90,90.1,90.5,null,11]
["2011-10-12",47063242,93.2,93.6,93.4,93.6,null,12]
["2011-10-13",47063363,81.2,79.3,78.7,78.900000000000006]
["2011-10-13",47063317,87.2,85.9,87.1,87.8]
["2011-10-13",47063273,93.5,94.2,93.8,94.8]
["2011-10-13",47063139,93.7,94.5,94,95.4]
["2011-10-14",47063445,89.8,89,88.5,88.8]
["2011-10-14",47063456,94.3,92.5,92.3,92.6]
["2011-10-14",47063399,93.4,92.6,92.4,92.6]
["2011-10-14",47063598,96,97.1,97.1,97.5]
["2011-10-14",47063611,92.4,92,92.5,92.6]
["2011-10-15",47063791,94.3,93.8,95.2,95.5]
["2011-10-16",47063944,95.2,95.9,95.5,96]
["2011-10-16",47063949,94.4,94.8,94.2,94.3]
["2011-10-16",47063979,83.5,82.7,82.1,81.7]
["2011-11-18",47064503,74.099999999999994,74.900000000000006,71.5,71.8]
["2011-10-19",47064495,96.3,96.4,96.7,97.9]
["2011-10-19",47064940,91,90,90.6,90.5]
["2011-10-20",47065180,94.8,95.6,95.2,95.2]
["2011-10-20",47065073,90.5,90.3,88.7,88.6]
["2011-10-21",47065432,96.8,98,96.9,98]
["2011-10-21",47065421,95.1,95.2,94.9,95]
["2011-10-26",47066431,95.3,95.9,95.6,96]
["2011-10-26",47066582,90.3,91.6,91.1,90.9]
["2011-10-27",47066702,87.2,85.4,85.7,86.3]
["2011-10-27",47066712,96.5,96.6,96.5,96.7]
["2011-10-28",47066955,97,97.6,97.4,97.4]
["2011-11-01",47067763,91.2,90.9,90.5,91.2]
["2011-11-01",47067808,94.4,94.7,94.1,94.5]
["2011-11-04",47068634,95.8,96.2,96,95.5]

{"sheet":"Na","index":43,"n_rows":44,"n_cols":6,"rows_kept":44,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Innkjøring Natrium",null,"Ref. området:  137 - 145 mmol/L   (7-350)"]
[]
["Dato","Labnummer","735 MB","835 plab","835 MB","837 HOV"]
["2011-10-10",47062477,140,141,139,142]
["2011-10-10",47062478,129,129,127,130]
["2011-10-10",47062589,132,132,131,133]
["2011-10-10",47062611,138,137,136,139]
["2011-10-10",47062649,132,133,131,132]
["2011-10-10",47062666,138,139,137,140]
["2011-10-11",47062501,134,134,133,136]
["2011-10-11",47062700,133,133,131,134]
["2011-10-11",47062193,136,138,140,137]
["2011-10-11",47062839,137,136,135,136]
["2011-10-12",47063235,138,139,139,140]
["2011-10-12",47063242,136,136,137,137]
["2011-10-13",47063363,138,138,139,139]
["2011-10-13",47063317,139,140,139,140]
["2011-10-13",47063273,135,135,136,137]
["2011-10-13",47063139,142,141,143,144]
["2011-10-14",47063445,142,142,143,144]
["2011-10-14",47063456,136,135,136,137]
["2011-10-14",47063399,145,145,145,146]
["2011-10-14",47063598,146,146,148,149]
["2011-10-14",47063611,144,144,145,146]
["2011-10-15",47063791,141,142,143,144]
["2011-10-16",47063944,139,142,141,141]
["2011-10-16",47063949,141,141,142,142]
["2011-10-16",47063979,140,140,141,141]
["2011-11-18",47064503,134,133,135,135]
["2011-10-19",47064495,146,146,146,147]
["2011-10-19",47064940,136,136,137,138]
["2011-10-20",47065180,137,139,139,139]
["2011-10-20",47065073,140,142,141,141]
["2011-10-21",47065432,141,141,143,143]
["2011-10-21",47065421,142,142,144,144]
["2011-10-26",47066431,139,141,141,142]
["2011-10-26",47066582,137,142,140,140]
["2011-10-27",47066702,127,129,129,130]
["2011-10-27",47066712,135,137,136,137]
["2011-10-28",47066955,137,139,138,139]
["2011-11-01",47067763,129,131,130,134]
["2011-11-01",47067808,130,131,130,130]
["2011-11-04",47068634,138,138,138,138]

{"sheet":"Na Passing-Bablok 735-835plab","index":44,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-09T14:18:06.670"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",0,0,0,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",1,1,1,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","< 0.01 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":"(non-linear relationship between X and Y detected)"}]

{"sheet":"Na Passing-Bablok 735-835MB","index":45,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-09T14:21:03.730"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-13,-31.307692307692317,1,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",1.1000000000000001,1,1.2307692307692308,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"Na Passing-Bablok 735-835HOV","index":46,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-09T14:22:43.380"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",2,-15.6875,2,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",1,1,1.125,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"Na Bias plot 735-835plab","index":47,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-09T14:35:22.700"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",0.65]
[null,null,null,null,"95% CI ",0.24969797527445778,1.0503020247255423," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-1.8032194124770273,-2.4661255174367969,-1.140313307517258]
[null,null,null,null,"Upper ",3.1032194124770269,2.4403133075172576,3.7661255174367962]

{"sheet":"Na Bias plot735-835MB","index":48,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-09T14:37:08.160"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",0.625]
[null,null,null,null,"95% CI ",0.17452829980431908,1.0754717001956808," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-2.1356803149428631,-2.8816681484444948,-1.3896924814412317]
[null,null,null,null,"Upper ",3.3856803149428627,2.6396924814412315,4.1316681484444944]

{"sheet":"Na Bias plot 735-835HOV","index":49,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-09T14:38:08.690"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",1.6]
[null,null,null,null,"95% CI ",1.2618491945007693,1.9381508054992309," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.47233056331459639,-1.0323133251209982,8.7652198491805292E-2]
[null,null,null,null,"Upper ",3.672330563314596,3.1123478015081947,4.2323133251209972]

{"sheet":"Na Descriptives ","index":50,"n_rows":35,"n_cols":13,"rows_kept":35,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Comparative descriptives"]
[null,null," "]
[null,"Variables "," 735 MB, 835 plab, 835 MB, 837 HOV"]
[null,"Performed by "," tingstad",null,null,null,null,null,null,null,"Date ","2011-12-13T13:15:22.940"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null," ","n","Mean","SD","SE","95% CI of Mean",null,null,"Median","IQR","95% CI of Median"]
[null,"735 MB ",40,137.47499999999999,4.5459900616945514,0.71878414077220754,136.02112199253136,138.92887800746863,null,138,5.75,136,139]
[null,"835 plab ",40,138.125,4.4846604650942625,0.70908708011039601,136.69073614709265,139.55926385290735,null,139,6.75,136,141]
[null,"835 MB ",40,138.1,5.0677461737986942,0.80128102564036952,136.47925630951056,139.72074369048943,null,139,6,136,141]
[null,"837 HOV ",40,139.07499999999999,4.6815924312468402,0.74022475796726261,137.57775425521186,140.57224574478812,null,139,5.25,137,141]

{"sheet":"K","index":51,"n_rows":44,"n_cols":6,"rows_kept":44,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Innkjøring Kalium",null,"Ref. området:  3,5 - 4,4 mmol/L    (0,5-25,0)"]
[]
["Dato","Labnummer","735 MB","835 plab","835 MB","837 HOV"]
["2011-10-10",47062477,3.5,3.6,3.5,3.5]
["2011-10-10",47062478,3.3,3.3,3.3,3.3]
["2011-10-10",47062589,4,4,3.9,3.9]
["2011-10-10",47062611,3.7,3.7,3.6,3.6]
["2011-10-10",47062649,3.7,3.8,3.7,3.8]
["2011-10-10",47062666,4.2,4.2,4.2,4.2]
["2011-10-11",47062501,3.6,3.7,3.6,3.6]
["2011-10-11",47062700,3.9,3.9,3.8,3.8]
["2011-10-11",47062193,4.7,4.8,4.9000000000000004,4.7]
["2011-10-11",47062839,4,4,4,3.9]
["2011-10-12",47063235,3.4,3.5,3.5,3.5]
["2011-10-12",47063242,3.8,3.8,3.8,3.8]
["2011-10-13",47063363,3,3,3,2.9]
["2011-10-13",47063317,4.4000000000000004,4.5,4.5,4.4000000000000004]
["2011-10-13",47063273,4.8,4.8,4.9000000000000004,4.8]
["2011-10-13",47063139,3.8,3.9,3.9,3.9]
["2011-10-14",47063445,4.5999999999999996,4.7,4.7,4.7]
["2011-10-14",47063456,3.9,4,4,3.9]
["2011-10-14",47063399,4.7,4.8,4.8,4.7]
["2011-10-14",47063598,4.0999999999999996,4.2,4.2,4.2]
["2011-10-14",47063611,2.6,2.6,2.6,2.6]
["2011-10-15",47063791,6.5,6.6,6.6,6.5]
["2011-10-16",47063944,3.7,3.8,3.7,3.6]
["2011-10-16",47063949,3.9,3.9,3.9,3.8]
["2011-10-16",47063979,3.7,3.8,3.8,3.7]
["2011-11-18",47064503,4.4000000000000004,4.5,4.5,4.4000000000000004]
["2011-10-19",47064495,4.7,4.7,4.8,4.7]
["2011-10-19",47064940,3.7,3.8,3.8,3.8]
["2011-10-20",47065180,3.6,3.7,3.7,3.6]
["2011-10-20",47065073,3.9,4,4,3.9]
["2011-10-21",47065432,3.5,3.6,3.6,3.6]
["2011-10-21",47065421,3.6,3.7,3.7,3.6]
["2011-10-26",47066431,3.7,3.8,3.8,3.7]
["2011-10-26",47066582,4.4000000000000004,4.5999999999999996,4.5,4.4000000000000004]
["2011-10-27",47066702,4.4000000000000004,4.5999999999999996,4.5999999999999996,4.5]
["2011-10-27",47066712,3.6,3.7,3.7,3.6]
["2011-10-28",47066955,3.9,4,4,3.9]
["2011-11-01",47067763,3.8,3.9,3.9,4.0999999999999996]
["2011-11-01",47067808,4.0999999999999996,4.2,4.2,4.0999999999999996]
["2011-11-04",47068634,3.8,3.8,3.8,3.8]

{"sheet":"K Passing-Bablok 735-835plab","index":52,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-09T14:39:25.410"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",9.9999999999999645E-2,9.9999999999997174E-2,9.9999999999999645E-2,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":"(constant bias detected)"}]
[null,"Slope ",1,1,1,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"K Passing-Bablok 735-835MB","index":53,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-09T14:41:42.640"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",9.9999999999995731E-2,-0.23846153846154058,9.9999999999999645E-2,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",1,1,1.0769230769230773,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"K Bias plot 735-835plab","index":54,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-09T14:44:31.100"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",7.2499999999999939E-2]
[null,null,null,null,"95% CI ",5.477831785988712E-2,9.0221682140112766E-2," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-3.6105932427346506E-2,-6.5453301551676848E-2,-6.7585633030161603E-3]
[null,null,null,null,"Upper ",0.18110593242734635,0.15175856330301601,0.21045330155167669]

{"sheet":"K Bias plot 735-835MB","index":55,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-09T14:50:47.950"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",0.06]
[null,null,null,null,"95% CI ",3.7327718908689196E-2,8.2672281091310765E-2," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-7.8945287964691316E-2,-0.11649092256171477,-4.1399653367667856E-2]
[null,null,null,null,"Upper ",0.19894528796469124,0.16139965336766779,0.23649092256171469]

{"sheet":"K Passing-Bablok 735-835HOV","index":56,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-09T14:43:15.400"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",0,0,0,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",1,1,1,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"K Bias plot 725-835HOV","index":57,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-09T14:52:14.950"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",9.9999999999999863E-3]
[null,null,null,null,"95% CI ",-1.4878442214143405E-2,3.4878442214143379E-2," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.14246557254805264,-0.18366464276747485,-0.10126650232863045]
[null,null,null,null,"Upper ",0.16246557254805258,0.1212665023286304,0.20366464276747476]

{"sheet":"K Descriptives  ","index":58,"n_rows":35,"n_cols":13,"rows_kept":35,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Comparative descriptives"]
[null,null," "]
[null,"Variables "," 735 MB, 835 plab, 835 MB, 837 HOV"]
[null,"Performed by "," tingstad",null,null,null,null,null,null,null,"Date ","2011-12-13T13:17:12.720"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null," ","n","Mean","SD","SE","95% CI of Mean",null,null,"Median","IQR","95% CI of Median"]
[null,"735 MB ",40,3.9649999999999999,0.62163968457290231,9.8289864359950338E-2,3.7661900040024783,4.1638099959975232,null,3.85,0.67500000000000004,3.7,4]
[null,"835 plab ",40,4.0374999999999996,0.63516099311866359,0.10042777095747558,3.8343656801786041,4.2406343198213969,null,3.9,0.72499999999999998,3.8,4]
[null,"835 MB ",40,4.0250000000000004,0.65073922462192069,0.10289090563085963,3.8168835205455065,4.2331164794544955,null,3.9,0.72499999999999998,3.8,4]
[null,"837 HOV ",40,3.9750000000000001,0.62869421045243468,9.9405282839546583E-2,3.7739338574007077,4.176066142599292,null,3.85,0.75,3.7,4.0999999999999996]

{"sheet":"Ca","index":59,"n_rows":44,"n_cols":6,"rows_kept":44,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Innkjøring Kalsium",null,"Ref. området:    1,15 - 1,29 ??mmol/L  (0,20-9,99)"]
[]
["Dato","Labnummer","735 MB","835 plab","835 MB","837 HOV"]
["2011-10-10",47062477,1.18,1.17,1.1399999999999999,1.1499999999999999]
["2011-10-10",47062478,1.1599999999999999,1.17,1.1499999999999999,1.1599999999999999]
["2011-10-10",47062589,1.19,1.18,1.1599999999999999,1.17]
["2011-10-10",47062611,1.18,1.1599999999999999,1.1399999999999999,1.1599999999999999]
["2011-10-10",47062649,1.19,1.2,1.17,1.2]
["2011-10-10",47062666,1.19,1.21,1.18,1.21]
["2011-10-11",47062501,1.1499999999999999,1.1200000000000001,1.1200000000000001,1.1200000000000001]
["2011-10-11",47062700,1,0.98,0.97,0.98]
["2011-10-11",47062193,1.3,1.3,1.35,1.29]
["2011-10-11",47062839,1.23,1.21,1.19,1.18]
["2011-10-12",47063235,1.24,1.24,1.24,1.24]
["2011-10-12",47063242,1.28,1.27,1.27,1.27]
["2011-10-13",47063363,1.1000000000000001,1.1000000000000001,1.1000000000000001,1.0900000000000001]
["2011-10-13",47063317,1.08,1.1000000000000001,1.0900000000000001,1.08]
["2011-10-13",47063273,1.1599999999999999,1.17,1.17,1.18]
["2011-10-13",47063139,1.1299999999999999,1.1299999999999999,1.1399999999999999,1.1399999999999999]
["2011-10-14",47063445,1.22,1.21,1.22,1.22]
["2011-10-14",47063456,1.17,1.1599999999999999,1.17,1.1599999999999999]
["2011-10-14",47063399,1.18,1.1599999999999999,1.17,1.1599999999999999]
["2011-10-14",47063598,1.26,1.24,1.26,1.25]
["2011-10-14",47063611,1.25,1.22,1.22,1.22]
["2011-10-15",47063791,1.18,1.1599999999999999,1.17,1.17]
["2011-10-16",47063944,1.24,1.23,1.25,1.23]
["2011-10-16",47063949,1.22,1.21,1.22,1.21]
["2011-10-16",47063979,1.21,1.18,1.2,1.19]
["2011-11-18",47064503,1.17,1.1100000000000001,1.1399999999999999,1.1399999999999999]
["2011-10-19",47064495,1.28,1.21,1.27,1.26]
["2011-10-19",47064940,1.22,1.24,1.23,1.23]
["2011-10-20",47065180,1.1299999999999999,1.1100000000000001,1.1299999999999999,1.1200000000000001]
["2011-10-20",47065073,1.25,1.26,1.26,1.25]
["2011-10-21",47065432,1.21,1.21,1.23,1.21]
["2011-10-21",47065421,1.21,1.2,1.22,1.21]
["2011-10-26",47066431,1.22,1.2,1.21,1.19]
["2011-10-26",47066582,1.25,1.25,1.25,1.23]
["2011-10-27",47066702,1.1200000000000001,1.1100000000000001,1.1100000000000001,1.1100000000000001]
["2011-10-27",47066712,1.2,1.2,1.2,1.19]
["2011-10-28",47066955,1.19,1.23,1.22,1.22]
["2011-11-01",47067763,1.18,1.1599999999999999,1.17,1.18]
["2011-11-01",47067808,1.18,1.18,1.17,1.17]
["2011-11-04",47068634,1.1299999999999999,1.08,1.1000000000000001,1.1000000000000001]

{"sheet":"Ca Passing-Bablok725-835plab","index":60,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-13T14:03:01.030"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-0.01,-0.20833333333333343,-9.9999999999946139E-3,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":"(constant bias detected)"}]
[null,"Slope ",1,0.99999999999999556,1.1666666666666667,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"Ca Passing-Bablok 725-735MB","index":61,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-13T14:06:51.450"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-0.13944444444444451,-0.30499999999999999,-0.01,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":"(constant bias detected)"}]
[null,"Slope ",1.1111111111111112,1,1.25,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","0.05 < p < 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":"(non-linear relationship between X and Y detected)"}]

{"sheet":"Ca Passing-Bablok 725-837HOV","index":62,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-13T14:13:20.670"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-0.01,-0.15749999999999686,-9.9999999999970529E-3,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":"(constant bias detected)"}]
[null,"Slope ",1,0.99999999999999756,1.125,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"Ca Bias plot 725-835plab","index":63,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-13T14:23:08.790"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-9.9999999999999863E-3]
[null,null,null,null,"95% CI ",-1.6755255905351932E-2,-3.2447440946480419E-3," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-5.1399053463748408E-2,-6.2585857685619736E-2,-4.0212249241877079E-2]
[null,null,null,null,"Upper ",3.1399053463748425E-2,2.0212249241877096E-2,4.2585857685619753E-2]

{"sheet":"Ca Bias plot 725-835MB","index":64,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-13T14:25:41.080"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-6.4999999999999945E-3]
[null,null,null,null,"95% CI ",-1.2668802808887639E-2,-3.311971911123494E-4," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-4.4305021877872997E-2,-5.4520651052140483E-2,-3.408939270360551E-2]
[null,null,null,null,"Upper ",3.1305021877872999E-2,2.1089392703605516E-2,4.1520651052140478E-2]

{"sheet":"Ca Bias plot 725-837HOV","index":65,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-13T14:27:32.460"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-9.7499999999999913E-3]
[null,null,null,null,"95% CI ",-1.4896051315109805E-2,-4.6039486848901776E-3," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-4.1287169946831889E-2,-4.9809107440306866E-2,-3.2765232453356913E-2]
[null,null,null,null,"Upper ",2.17871699468319E-2,1.3265232453356927E-2,3.0309107440306873E-2]

{"sheet":"Ca Descriptives ","index":66,"n_rows":35,"n_cols":13,"rows_kept":35,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Comparative descriptives"]
[null,null," "]
[null,"Variables "," 735 MB, 835 plab, 835 MB, 837 HOV"]
[null,"Performed by "," tingstad",null,null,null,null,null,null,null,"Date ","2011-12-13T14:16:02.850"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null," ","n","Mean","SD","SE","95% CI of Mean",null,null,"Median","IQR","95% CI of Median"]
[null,"735 MB ",40,1.19075,5.8238567712734571E-2,9.208326081909201E-3,1.172124404335487,1.2093755956645129,null,1.19,6.0000000000000053E-2,1.18,1.22]
[null,"835 plab ",40,1.1807500000000002,6.0780711346568414E-2,9.6102742830198006E-3,1.1613113874412797,1.2001886125587207,null,1.19,5.7500000000000107E-2,1.1599999999999999,1.21]
[null,"835 MB ",40,1.18425,6.5979309111032919E-2,1.0432244761758171E-2,1.163148795386115,1.205351204613885,null,1.175,8.7500000000000133E-2,1.17,1.22]
[null,"837 HOV ",40,1.181,5.9734455117671628E-2,9.4448466480471852E-3,1.1618959963825062,1.2001040036174939,null,1.1850000000000001,6.25E-2,1.1599999999999999,1.21]

{"sheet":"Cl","index":67,"n_rows":44,"n_cols":5,"rows_kept":44,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Innkjøring Klorid"]
[]
["Dato","Labnummer","725 MB","835 MB","837 HOV"]
["2011-10-10",47062477,99,96,100]
["2011-10-10",47062478,98,97,96]
["2011-10-10",47062589,100,98,100]
["2011-10-10",47062611,102,102,103]
["2011-10-10",47062649,103,99,102]
["2011-10-10",47062666,104,101,103]
["2011-10-11",47062501,107,104,106]
["2011-10-11",47062700,104,102,104]
["2011-10-11",47062193,104,102,103]
["2011-10-11",47062839,100,98,100]
["2011-10-12",47063235,114,111,112]
["2011-10-12",47063242,101,100,101]
["2011-10-13",47063363,88,87,87]
["2011-10-13",47063317,105,104,105]
["2011-10-13",47063273,107,107,106]
["2011-10-13",47063139,109,109,108]
["2011-10-14",47063445,113,112,111]
["2011-10-14",47063456,108,106,106]
["2011-10-14",47063399,113,112,111]
["2011-10-14",47063598,117,118,119]
["2011-10-14",47063611,111,111,112]
["2011-10-15",47063791,116,114,114]
["2011-10-16",47063944,111,108,109]
["2011-10-16",47063949,110,109,108]
["2011-10-16",47063979,109,108,108]
["2011-11-18",47064503,103,102,101]
["2011-10-19",47064495,114,114,114]
["2011-10-19",47064940,106,105,104]
["2011-10-20",47065180,106,105,104]
["2011-10-20",47065073,102,101,99]
["2011-10-21",47065432,112,111,110]
["2011-10-21",47065421,108,106,107]
["2011-10-26",47066431,109,106,107]
["2011-10-26",47066582,105,106,105]
["2011-10-27",47066702,98,95,96]
["2011-10-27",47066712,104,104,105]
["2011-10-28",47066955,111,107,108]
["2011-11-01",47067763,101,98,102]
["2011-11-01",47067808,98,97,98]
["2011-11-04",47068634,102,102,102]

{"sheet":"Cl Passing-Bablok725-825MB","index":68,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-13T13:20:48.050"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-1,-14.5625,-1,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":"(constant bias detected)"}]
[null,"Slope ",1,1,1.125,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"Cl Passing-Bablok 725-837HOV","index":69,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-13T13:30:12.430"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-1,-1,14.28571428571429,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",1,0.8571428571428571,1,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","< 0.01 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":"(non-linear relationship between X and Y detected)"}]

{"sheet":"Cl Bias plot 725-835MB","index":70,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 725 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-13T13:39:30.200"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-1.45]
[null,null,null,null,"95% CI ",-1.8593715060891549,-1.040628493910845," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-3.9588010142877308,-4.6367263152380271,-3.2808757133374344]
[null,null,null,null,"Upper ",1.0588010142877304,0.38087571333743403,1.7367263152380268]

{"sheet":"Cl Bias plot 725-837HOV","index":71,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 725 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-13T13:35:21.010"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-0.9]
[null,null,null,null,"95% CI ",-1.2953578001094752,-0.50464219989052472," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-3.3229191215501994,-3.9776375170310314,-2.6682007260693674]
[null,null,null,null,"Upper ",1.5229191215501987,0.86820072606936693,2.1776375170310303]

{"sheet":"Cl Descriptives ","index":72,"n_rows":34,"n_cols":13,"rows_kept":34,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Comparative descriptives"]
[null,null," "]
[null,"Variables "," 725 MB, 835 MB, 837 HOV"]
[null,"Performed by "," tingstad",null,null,null,null,null,null,null,"Date ","2011-12-13T13:37:06.940"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null," ","n","Mean","SD","SE","95% CI of Mean",null,null,"Median","IQR","95% CI of Median"]
[null,"725 MB ",40,105.8,5.949358939824676,0.94067624338653033,103.89730289689113,107.70269710310886,null,105.5,8.75,103,109]
[null,"835 MB ",40,104.35,6.2081418831261921,0.98159341940828049,102.36454010487705,106.33545989512294,null,104.5,8,102,107]
[null,"837 HOV ",40,104.9,5.8520213470350457,0.92528581862786963,103.02843296614695,106.77156703385306,null,105,6.25,103,107]

{"sheet":"Glukose","index":73,"n_rows":43,"n_cols":6,"rows_kept":43,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Innkjøring Glukose",null,"Ref. området:  4,0 - 6,0 mmol/L    (0,0-60)"]
[]
["Dato","Labnummer","735 MB","835 plab","835 MB","837 HOV"]
["2011-10-10",47062477,6.9,6.7,6.7,6.6]
["2011-10-10",47062478,5.9,6,6,5.8]
["2011-10-10",47062589,5.8,5.8,5.7,5.6]
["2011-10-10",47062611,10,10,9.9,9.8000000000000007]
["2011-10-10",47062666,6.3,6.3,6.3,6.3]
["2011-10-11",47062501,10,9.6999999999999993,9.8000000000000007,9.4]
["2011-10-11",47062700,7.8,7.5,7.4,7.3]
["2011-10-11",47062193,8.4,8,8.4,8.3000000000000007]
["2011-10-11",47062839,19.100000000000001,18.7,18.3,18.2]
["2011-10-12",47063235,4.8,4.8,4.9000000000000004,4.8]
["2011-10-12",47063242,8,7.7,8,7.8]
["2011-10-13",47063363,6.4,6.3,6.4,6.2]
["2011-10-13",47063317,5.9,5.8,6,5.7]
["2011-10-13",47063273,8.1999999999999993,8.1,8.1999999999999993,8.1]
["2011-10-13",47063139,10.5,9.9,10.199999999999999,10]
["2011-10-14",47063445,6.8,6.5,6.6,6.6]
["2011-10-14",47063456,6.1,5.8,6,5.7]
["2011-10-14",47063399,9.6,9.1999999999999993,9.5,9.4]
["2011-10-14",47063598,6,5.7,5.8,5.6]
["2011-10-14",47063611,2.4,2.1,2.2000000000000002,2.1]
["2011-10-15",47063791,13.6,13.1,13,13.1]
["2011-10-16",47063944,5.6,5.6,5.7,5.5]
["2011-10-16",47063949,6.9,6.7,6.8,6.7]
["2011-10-16",47063979,7.5,7.2,7.3,7.1]
["2011-11-18",47064503,10,9.9,10.1,10]
["2011-10-19",47064495,8.8000000000000007,8.6,8.6999999999999993,8.5]
["2011-10-19",47064940,7.5,7.3,7.4,7.3]
["2011-10-20",47065180,9.8000000000000007,9.5,9.6,9.5]
["2011-10-20",47065073,8,7.8,8,8]
["2011-10-21",47065432,3.8,3.7,3.8,3.7]
["2011-10-21",47065421,7.6,7.4,7.4,7.3]
["2011-10-26",47066431,5.7,5.6,5.8,5.6]
["2011-10-26",47066582,6.1,5.8,5.9,5.8]
["2011-10-27",47066702,8.1,8,8.1,8]
["2011-10-27",47066712,8.1999999999999993,8,8.1,8.1]
["2011-10-28",47066955,6,5.6,5.9,5.7]
["2011-11-01",47067763,6.2,6.1,6.1,5.8]
["2011-11-01",47067808,5,5,5,4.9000000000000004]
["2011-11-04",47068634,6.7,6.4,6.6,6.6]

{"sheet":"Glu Passing-Bablok725-835plab","index":74,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-13T14:29:55.980"]
[]
[]
[null,"n ",39," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",2.7027027027037853E-3,-0.2,0.22777777777777347,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.9729729729729728,0.94444444444444509,1,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"Glu Passing-Bablok 725-835MB","index":75,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-13T14:32:36.270"]
[]
[]
[null,"n ",39," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",0.11515151515151505,-9.9999999999999645E-2,0.3,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.96969696969696972,0.94666666666666666,1,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"Glu Passing-Bablok 725-837HOV","index":76,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-13T14:39:01.090"]
[]
[]
[null,"n ",39," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-4.5454545454545019E-2,-0.2,0.15844155844155763,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.97727272727272718,0.94805194805194815,1,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"Glu Bias plot 725-835plab","index":77,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-13T14:41:18.090"]
[]
[]
[null,null,null,null,"n ",39," "]
[]
[null,null,null,null,"Bias ",-0.20769230769230779]
[null,null,null,null,"95% CI ",-0.25834154503999585,-0.15704307034461973," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.51392973700572842,-0.59773505698410767,-0.43012441702734916]
[null,null,null,null,"Upper ",9.854512162111273E-2,1.4739801642733487E-2,0.18235044159949199]

{"sheet":"GluBias plot 725-837HOV","index":78,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-13T14:44:02.530"]
[]
[]
[null,null,null,null,"n ",39," "]
[]
[null,null,null,null,"Bias ",-0.24358974358974372]
[null,null,null,null,"95% CI ",-0.30436929504639193,-0.18281019213309552," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.6110774801796478,-0.71164463626535102,-0.51051032409394459]
[null,null,null,null,"Upper ",0.12389799300016026,2.3330836914457115E-2,0.2244651490858634]

{"sheet":"Glu Bias plot 725-835MB","index":79,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-13T14:42:32.490"]
[]
[]
[null,null,null,null,"n ",39," "]
[]
[null,null,null,null,"Bias ",-0.11282051282051285]
[null,null,null,null,"95% CI ",-0.17239825650983864,-5.3242769131187057E-2," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.47304183158328805,-0.57162045062939937,-0.37446321253717674]
[null,null,null,null,"Upper ",0.24740080594226227,0.14882218689615098,0.34597942498837359]

{"sheet":"Glu Descriptives ","index":80,"n_rows":35,"n_cols":13,"rows_kept":35,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Comparative descriptives"]
[null,null," "]
[null,"Variables "," 735 MB, 835 plab, 835 MB, 837 HOV"]
[null,"Performed by "," tingstad",null,null,null,null,null,null,null,"Date ","2011-12-13T14:45:09.970"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null," ","n","Mean","SD","SE","95% CI of Mean",null,null,"Median","IQR","95% CI of Median"]
[null,"735 MB ",39,7.5897435897435912,2.791226373722798,0.44695392607629986,6.684932758180854,8.4945544213063275,null,6.9,2.2999999999999998,6.1,8.1]
[null,"835 plab ",39,7.3820512820512816,2.7226991053646703,0.43598078110880706,6.4994544189618875,8.2646481451406757,null,6.7,2.25,6,8]
[null,"835 MB ",39,7.4769230769230761,2.6689699904113944,0.42737723712575776,6.6117431763694592,8.3421029774766939,null,6.8,2.35,6,8.1]
[null,"837 HOV ",39,7.3461538461538476,2.6864030102728886,0.43016875441142616,6.4753228148478046,8.2169848774598915,null,6.7,2.5,5.8,8]

{"sheet":"Lakt","index":81,"n_rows":42,"n_cols":6,"rows_kept":42,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Innkjøring Laktat",null,"Ref. området:  0,50 - 1,60 mmol/L    (0,0-30)"]
[]
["Dato","Labnummer","735 MB","835 plab","835 MB","837 HOV"]
["2011-10-10",47062477,0.9,1,1,0.9]
["2011-10-10",47062478,1.8,1.8,1.8,1.8]
["2011-10-10",47062589,0.8,0.8,0.8,0.7]
["2011-10-10",47062611,1.9,1.9,1.7,1.7]
["2011-10-10",47062649,1.2,1.2,1.2,0.5]
["2011-10-10",47062666,1,1,0.9,1]
["2011-10-11",47062501,1.6,1.7,1.6,1.4]
["2011-10-11",47062700,0.9,1,0.9,0.8]
["2011-10-11",47062193,1.4,1.3,1.2,1.3]
["2011-10-11",47062839,1.5,1.4,1.3,1.3]
["2011-10-12",47063235,0.5,0.5,0.5,0.4]
["2011-10-12",47063242,1.1000000000000001,1,1,0.9]
["2011-10-13",47063363,1.2,1.1000000000000001,1.1000000000000001,1.1000000000000001]
["2011-10-13",47063317,0.9,0.9,0.9,0.8]
["2011-10-13",47063273,1.3,1.3,1.2,1.1000000000000001]
["2011-10-13",47063139,1.4,1.3,1.4,1.3]
["2011-10-14",47063456,0.8,0.7,0.7,0.6]
["2011-10-14",47063399,1.3,1.2,1.2,1.1000000000000001]
["2011-10-14",47063611,1.1000000000000001,1.3,1.2,1.1000000000000001]
["2011-10-15",47063791,1.1000000000000001,1,1,0.8]
["2011-10-16",47063944,0.9,0.9,1,0.8]
["2011-10-16",47063949,1.4,1.3,1.4,1.2]
["2011-10-16",47063979,0.8,0.8,0.8,0.7]
["2011-11-18",47064503,1.9,1.8,1.8,1.6]
["2011-10-19",47064495,1.3,1.2,1.1000000000000001,1.1000000000000001]
["2011-10-19",47064940,1,0.9,0.9,0.9]
["2011-10-20",47065180,1.3,1.4,1.4,1.3]
["2011-10-20",47065073,1.5,1.4,1.4,1.3]
["2011-10-21",47065432,0.9,0.8,0.8,0.7]
["2011-10-21",47065421,2,1.9,1.9,1.8]
["2011-10-26",47066431,0.9,0.9,1,0.9]
["2011-10-26",47066582,1.5,1.4,1.4,1.3]
["2011-10-27",47066702,1.3,1.2,1.2,1.1000000000000001]
["2011-10-27",47066712,2.9,2.7,2.6,2.6]
["2011-10-28",47066955,0.7,0.8,0.8,0.7]
["2011-11-01",47067763,0.9,0.9,0.9,0.9]
["2011-11-01",47067808,0.6,0.6,0.6,0.8]
["2011-11-04",47068634,1.1000000000000001,1.1000000000000001,1.1000000000000001,1.5]

{"sheet":"Lakt Passing-Bablok 725-835plab","index":82,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-13T14:54:35.090"]
[]
[]
[null,"n ",38," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",1.2213537473049207E-16,0,8.5000000000000075E-2,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",1,0.9,1,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","0.01 < p < 0.05 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":"(non-linear relationship between X and Y detected)"}]

{"sheet":"Lakt Passing-Bablok 725-835MB","index":83,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-13T14:57:42.450"]
[]
[]
[null,"n ",38," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",5.416666666666678E-2,-4.9999999999999933E-2,0.15,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.91666666666666663,0.83333333333333337,1,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"Lakt Passing-Bablok 725-837HOV","index":84,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-13T14:59:47.070"]
[]
[]
[null,"n ",38," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-6.5694444444444527E-2,-0.1,0.05,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.9472222222222223,0.83333333333333337,1,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","< 0.01 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":"(non-linear relationship between X and Y detected)"}]

{"sheet":"Lakt Bias plot 725-835plab","index":85,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-13T15:02:00.030"]
[]
[]
[null,null,null,null,"n ",38," "]
[]
[null,null,null,null,"Bias ",-3.1578947368421061E-2]
[null,null,null,null,"95% CI ",-5.9243640914421192E-2,-3.9142538224209312E-3," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.19654139805966409,-0.24227537023894791,-0.15080742588038026]
[null,null,null,null,"Upper ",0.13338350332282195,8.7649531143538129E-2,0.17911747550210577]

{"sheet":"Lakt Bias plot 725-835MB","index":86,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-13T15:04:07.900"]
[]
[]
[null,null,null,null,"n ",38," "]
[]
[null,null,null,null,"Bias ",-0.05]
[null,null,null,null,"95% CI ",-8.2196018578736291E-2,-1.7803981421263725E-2," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.24198239512098388,-0.29520734104403473,-0.18875744919793302]
[null,null,null,null,"Upper ",0.14198239512098382,8.8757449197932978E-2,0.19520734104403464]

{"sheet":"Lakt Bias plot 725-837HOV","index":87,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-13T15:05:23.120"]
[]
[]
[null,null,null,null,"n ",38," "]
[]
[null,null,null,null,"Bias ",-0.12631578947368419]
[null,null,null,null,"95% CI ",-0.18072085450045858,-7.191072444690981E-2," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.45072902911106133,-0.54066892763946472,-0.36078913058265794]
[null,null,null,null,"Upper ",0.19809745016369287,0.10815755163528953,0.28803734869209624]

{"sheet":"Lakt Descriptives ","index":88,"n_rows":35,"n_cols":13,"rows_kept":35,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Comparative descriptives"]
[null,null," "]
[null,"Variables "," 735 MB, 835 plab, 835 MB, 837 HOV"]
[null,"Performed by "," tingstad",null,null,null,null,null,null,null,"Date ","2011-12-13T15:06:37.900"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null," ","n","Mean","SD","SE","95% CI of Mean",null,null,"Median","IQR","95% CI of Median"]
[null,"735 MB ",38,1.226315789473684,0.45836619995094313,7.4356816354349251E-2,1.0756545833114646,1.3769769956359035,null,1.1499999999999999,0.5,0.9,1.3]
[null,"835 plab ",38,1.1947368421052631,0.42929695412638397,6.964116198548978E-2,1.0536304583505922,1.335843225859934,null,1.1499999999999999,0.42499999999999999,1,1.3]
[null,"835 MB ",38,1.1763157894736842,0.4089764484901669,6.6344740683086989E-2,1.0418885890647887,1.3107429898825798,null,1.1000000000000001,0.5,1,1.2]
[null,"837 HOV ",38,1.1000000000000001,0.42680459482316047,6.9236847917352379E-2,0.95971283428290177,1.2402871657170988,null,1.1000000000000001,0.5,0.9,1.3]

{"sheet":"Kreat","index":89,"n_rows":38,"n_cols":4,"rows_kept":38,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Innkjøring Kreatinin",null,"Ref. området:  50 - 100 umol/L  (10-1800)"]
[]
["Dato","Labnummer","837 HOV","Serum"]
["2011-10-10",47062478,42,54]
[null,47062477,46,58]
[null,47062589,71,91]
[null,47062611,60,69]
[null,47062649,51,62]
[null,47062666,62,72]
["2011-10-11",47062501,92,96]
[null,47062700,275,287]
[null,47062193,143,154]
["2011-10-12",47063235,38,44]
[null,47063242,58,65]
["1900-01-13",47063317,784,776]
[14,47063139,104,105]
["1900-01-15",47063445,147,156]
[16,47063456,112,122]
["1900-01-17",47063399,93,100]
[18,47063598,72,78]
["1900-01-19",47063611,69,74]
["2011-10-15",47063791,759,743]
["2011-10-16",47063944,69,68]
[null,47063949,60,69]
[null,47063979,54,61]
["2011-10-18",47064503,221,224]
["2011-10-19",47064495,87,89]
[null,47064837,47,52]
[null,47064940,77,85]
["2011-10-20",47065180,157,172]
[null,47065073,104,105]
["2011-10-21",47065432,67,67]
["2011-10-26",47066431,54,58]
[null,47066582,71,76]
["2011-10-28",47066955,74,88]
["2011-11-01",47067763,44,53]
[null,47067808,58,66]

{"sheet":"KreatPassing-Bablok Arch-837HOV","index":90,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," Serum  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-13T15:15:25.690"]
[]
[]
[null,"n ",34," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-8.5778688524590088,-11.320437956204385,-5.4274809160305262,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":"(constant bias detected)"}]
[null,"Slope ",1.0191256830601092,0.97709923664122134,1.0408759124087592,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"Kreat Bias plot Arch-837HOV","index":91,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," Serum  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-13T15:17:52.090"]
[]
[]
[null,null,null,null,"n ",34," "]
[]
[null,null,null,null,"Bias ",-6.382352941176471]
[null,null,null,null,"95% CI ",-8.6744023121880289,-4.090303570164914," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-19.257452640409632,-23.031060465896768,-15.483844814922497]
[null,null,null,null,"Upper ",6.4927467580566862,2.7191389325695519,10.266354583543821]

{"sheet":"Kreat Descriptives ","index":92,"n_rows":33,"n_cols":13,"rows_kept":33,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Comparative descriptives"]
[null,null," "]
[null,"Variables "," 837 HOV, Serum"]
[null,"Performed by "," tingstad",null,null,null,null,null,null,null,"Date ","2011-12-13T15:19:50.470"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null," ","n","Mean","SD","SE","95% CI of Mean",null,null,"Median","IQR","95% CI of Median"]
[null,"837 HOV ",34,127.11764705882354,171.2451308232136,29.368297648724958,67.367402103147725,186.86789201449935,null,71,46,60,93]
[null,"Serum ",34,133.5,166.98434874943439,28.637579551471262,75.236412045239916,191.76358795476008,null,77,39.75,67,100]

{"sheet":"sO2","index":93,"n_rows":43,"n_cols":6,"rows_kept":43,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Innkjøring sO2",null,"Ref. området:  95 - 98 %",null,"(0,0-100)"]
[]
["Dato","Labnummer","735 MB","835 plab","835 MB","837 HOV"]
["2011-10-10",47062477,89.9,88.2,88.4,89]
["2011-10-10",47062478,98,98,98.2,98.1]
["2011-10-10",47062589,97.3,97.6,97.8,97.6]
["2011-10-10",47062611,93.9,94.8,94.6,95.2]
["2011-10-10",47062649,98.2,98.5,98.3,98.4]
["2011-10-10",47062666,94.8,93.3,93.4,93.5]
["2011-10-11",47062501,98.4,97.6,97.7,97.5]
["2011-10-11",47062700,97.3,96.9,96.8,96.7]
["2011-10-11",47062193,94.2,93.8,94,93.1]
["2011-10-11",47062839,83.2,82.1,82,83.2]
["2011-10-12",47063242,95.7,95.2,95.4,95.5]
["2011-10-13",47063363,84.3,81.5,81.099999999999994,81.3]
["2011-10-13",47063317,89.3,87.6,88.9,88.9]
["2011-10-13",47063273,96.8,96.4,96.4,96.7]
["2011-10-13",47063139,96.2,96.4,96.5,97.1]
["2011-10-14",47063445,91.9,90.1,89.9,90]
["2011-10-14",47063456,97.5,95,94.9,95]
["2011-10-14",47063399,95.7,94.6,94.5,94.7]
["2011-10-14",47063598,98.6,98.1,98.4,98.2]
["2011-10-14",47063611,94,93.9,93.5,94.6]
["2011-10-15",47063791,97.6,97,97.3,97.7]
["2011-10-16",47063944,97.3,97.1,97.3,96.9]
["2011-10-16",47063949,96.7,96,95.5,95.2]
["2011-10-16",47063979,85.6,84.2,83.7,83.1]
["2011-11-18",47064503,76.2,76.900000000000006,73.400000000000006,73.7]
["2011-10-19",47064495,98.7,98.2,98.5,98.5]
["2011-10-19",47064940,93.9,91.6,92.4,92.3]
["2011-10-20",47065180,97.6,97.7,97.3,97.3]
["2011-10-20",47065073,93.3,91.8,91.3,90.6]
["2011-10-21",47065432,99.2,99.1,99.3,99.1]
["2011-10-21",47065421,97.8,96.5,96.2,96.4]
["2011-10-26",47066431,98.4,97.4,97.2,97.6]
["2011-10-26",47066582,97.4,97.8,97.7,97]
["2011-10-27",47066702,89.7,87.8,87.8,88.3]
["2011-10-27",47066712,98.9,98.4,98.9,98.6]
["2011-10-28",47066955,99.6,99.4,99.5,99.4]
["2011-11-01",47067763,94.8,93.7,93.7,94.1]
["2011-11-01",47067808,97.5,96.1,96.1,96]
["2011-11-04",47068634,98,97.8,97.7,97.3]

{"sheet":"sO2Passing-Bablok725-835plab","index":94,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-13T15:26:14.020"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-13.457831325301296,-19.850000000000264,-5.9805555555556458,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":"(constant bias detected)"}]
[null,"Slope ",1.1325301204819287,1.0555555555555565,1.2,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":"(proportional bias detected)"}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"sO2Passing-Bablok 725-835MB","index":95,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-13T15:38:25.040"]
[]
[]
[null,"n ",39," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-14.786658249158258,-22.823076923076666,-9.1999999999999833,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":"(constant bias detected)"}]
[null,"Slope ",1.1479377104377106,1.0909090909090908,1.2307692307692282,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":"(proportional bias detected)"}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"sO2Passing-Bablok725-837HOV","index":96,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-13T15:40:44.210"]
[]
[]
[null,"n ",39," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-13.681818181818047,-21.310734463276823,-6.9733333333332572,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":"(constant bias detected)"}]
[null,"Slope ",1.1363636363636349,1.066666666666666,1.214689265536723,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":"(proportional bias detected)"}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"sO2 Bias plot 725-835plab","index":97,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-13T15:42:52.220"]
[]
[]
[null,null,null,null,"n ",39," "]
[]
[null,null,null,null,"Bias ",-0.75128205128205228]
[null,null,null,null,"95% CI ",-1.0370788294814017,-0.46548527308270282," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-2.4792778759974992,-2.9521633862248247,-2.0063923657701737]
[null,null,null,null,"Upper ",0.97671377343339405,0.50382826320606877,1.4495992836607192]

{"sheet":"sO2 Bias plot 725-835MB","index":98,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-13T15:44:30.920"]
[]
[]
[null,null,null,null,"n ",39," "]
[]
[null,null,null,null,"Bias ",-0.81794871794871737]
[null,null,null,null,"95% CI ",-1.1277046439071345,-0.50819279199030032," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-2.6908072928751938,-3.2033361254801971,-2.1782784602701906]
[null,null,null,null,"Upper ",1.0549098569777584,0.54238102437275504,1.5674386895827617]

{"sheet":"sO2 Bias plot 725-837HOV","index":99,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-13T15:45:58.480"]
[]
[]
[null,null,null,null,"n ",39," "]
[]
[null,null,null,null,"Bias ",-0.76923076923077027]
[null,null,null,null,"95% CI ",-1.0934916527318539,-0.44496988572968665," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-2.7297897942089246,-3.2663188422523439,-2.1932607461655054]
[null,null,null,null,"Upper ",1.1913282557473834,0.65479920770396416,1.7278573037908027]

{"sheet":"sO2 Descriptives ","index":100,"n_rows":35,"n_cols":13,"rows_kept":35,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Comparative descriptives"]
[null,null," "]
[null,"Variables "," 735 MB, 835 plab, 835 MB, 837 HOV"]
[null,"Performed by "," tingstad",null,null,null,null,null,null,null,"Date ","2011-12-13T15:47:39.650"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null," ","n","Mean","SD","SE","95% CI of Mean",null,null,"Median","IQR","95% CI of Median"]
[null,"735 MB ",39,94.702564102564097,5.1077076524491867,0.81788779656280386,93.046836981360386,96.358291223767807,null,96.8,4,94.2,97.6]
[null,"835 plab ",39,93.95128205128205,5.3988727768358116,0.86451153038326267,92.201170124604133,95.701393977959967,null,96.1,5.1000000000000085,93.8,97.4]
[null,"835 MB ",39,93.884615384615401,5.7462820227535163,0.92014153154688216,92.021886419186828,95.747344350043974,null,96.1,4.8,93.7,97.3]
[null,"837 HOV ",39,93.933333333333323,5.6336551345852364,0.90210679587568354,92.107113777997469,95.759552888669177,null,96,4.8500000000000085,94.1,97.3]

{"sheet":"COHb","index":101,"n_rows":44,"n_cols":6,"rows_kept":44,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Innkjøring COHb",null,"Ref. området:  0,60 - 1,50 %      (0-100)"]
[]
["Dato","Labnummer","735 MB","835 plab","835 MB","837 HOV"]
["2011-10-10",47062477,3.1,3,2.9,2.7]
["2011-10-10",47062478,3.6,3.1,3.3,3]
["2011-10-10",47062589,1.2,1.3,1.4,0.9]
["2011-10-10",47062611,1.8,1.7,1.4,1.5]
["2011-10-10",47062649,1.7,1.3,1.4,0.9]
["2011-10-10",47062666,2.7,2.1,2.4,1.9]
["2011-10-11",47062501,2.4,1.5,1.7,1.3]
["2011-10-11",47062700,1.8,1.4,1.6,1.2]
["2011-10-11",47062193,1.4,1.1000000000000001,1.5,1.7]
["2011-10-11",47062839,1.6,1.4,1.6,1.3]
["2011-10-12",47063235,7.9,7.7,7.8,7.5]
["2011-10-12",47063242,1.4,1.2,1.6,1.3]
["2011-10-13",47063363,2.7,2.4,2.6,2.6]
["2011-10-13",47063317,1.2,1.2,1.4,0.9]
["2011-10-13",47063273,2.5,1.6,1.8,1.5]
["2011-10-13",47063139,1.7,1.4,1.7,1.3]
["2011-10-14",47063445,0.9,0.5,0.7,0.3]
["2011-10-14",47063456,2.2999999999999998,1.8,2,1.6]
["2011-10-14",47063399,1.4,1.3,1.5,1.1000000000000001]
["2011-10-14",47063598,1.1000000000000001,0.5,0.7,0.1]
["2011-10-14",47063611,0.6,1.1000000000000001,0.5,1.1000000000000001]
["2011-10-15",47063791,1.9,1.8,1.8,1.5]
["2011-10-16",47063944,1,0.9,1.1000000000000001,0.6]
["2011-10-16",47063949,1.4,1.1000000000000001,0.9,0.6]
["2011-10-16",47063979,1.5,1.4,1.5,1.2]
["2011-11-18",47064503,1.7,1.9,1.9,1.9]
["2011-10-19",47064495,1.3,1.1000000000000001,1.1000000000000001,0.8]
["2011-10-19",47064940,1.8,1.2,1.3,0.9]
["2011-10-20",47065180,2,1.8,1.6,1.5]
["2011-10-20",47065073,2,1.5,1.6,1.6]
["2011-10-21",47065432,1.3,1.1000000000000001,1.5,0.9]
["2011-10-21",47065421,1.5,0.9,0.8,1]
["2011-10-26",47066431,1.8,1.3,0.9,1]
["2011-10-26",47066582,5.8,5.5,5.7,5.5]
["2011-10-27",47066702,2,2.1,2,1.9]
["2011-10-27",47066712,1.2,1.1000000000000001,1.6,1.1000000000000001]
["2011-10-28",47066955,1.6,1.5,1.5,1.4]
["2011-11-01",47067763,2.9,2.8,2.8,2.7]
["2011-11-01",47067808,2.1,1.3,1.6,1.3]
["2011-11-04",47068634,1.3,1,1.2,1]

{"sheet":"COHbPassing-Bablok 725-835plab","index":102,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-14T08:29:55.040"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-0.24999999999999201,-0.3259615384615383,9.0000000000000246E-2,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.99999999999999778,0.8,1.0192307692307692,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"COHb Passing-Bablok 725-835MB","index":103,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-14T08:32:22.910"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-6.7792022792022788E-2,-0.22444444444444461,0.26,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.96609686609686607,0.8,1.0222222222222224,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"COHb Passing-Bablok 725-837HOV","index":104,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-14T08:34:10.730"]
[]
[]
[null,"n ",40," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-0.4,-0.63780487804878061,-0.15,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":"(constant bias detected)"}]
[null,"Slope ",1,0.875,1.1219512195121952,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"COHb Bias plot 725-835plab","index":105,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-14T08:36:14.670"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-0.28000000000000003]
[null,null,null,null,"95% CI ",-0.3719527449497646,-0.18804725505023523," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.84352515103058856,-0.99580026389518295,-0.69125003816599417]
[null,null,null,null,"Upper ",0.28352515103058862,0.13125003816599423,0.435800263895183]

{"sheet":"COHbBias plot 725-835MB","index":106,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-14T08:37:37.840"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-0.18]
[null,null,null,null,"95% CI ",-0.27392802166400188,-8.6071978335998264E-2," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.75563047871086408,-0.91117667918424694,-0.60008427823748123]
[null,null,null,null,"Upper ",0.39563047871086376,0.24008427823748091,0.55117667918424662]

{"sheet":"COHb Bias plot 725-837HOV","index":107,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-14T08:38:39.740"]
[]
[]
[null,null,null,null,"n ",40," "]
[]
[null,null,null,null,"Bias ",-0.42499999999999999]
[null,null,null,null,"95% CI ",-0.53549103056235381,-0.31450896943764672," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-1.1021355734860532,-1.2851103629673193,-0.91916078400478718]
[null,null,null,null,"Upper ",0.25213557348605259,6.9160784004786632E-2,0.43511036296731853]

{"sheet":"COHb Descriptives ","index":108,"n_rows":35,"n_cols":13,"rows_kept":35,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Comparative descriptives"]
[null,null," "]
[null,"Variables "," 735 MB, 835 plab, 835 MB, 837 HOV"]
[null,"Performed by "," tingstad",null,null,null,null,null,null,null,"Date ","2011-12-14T08:39:40.130"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null," ","n","Mean","SD","SE","95% CI of Mean",null,null,"Median","IQR","95% CI of Median"]
[null,"735 MB ",40,2.0274999999999999,1.3044278635792979,0.20624815461489901,1.6103237727206365,2.4446762272793641,null,1.7,0.875,1.4,2]
[null,"835 plab ",40,1.7475000000000001,1.2903959837475452,0.20402951960879306,1.3348113851469647,2.1601886148530349,null,1.4,0.7,1.2,1.6]
[null,"835 MB ",40,1.8474999999999999,1.3046244179053494,0.2062792325826131,1.4302609116039064,2.2647390883960932,null,1.6,0.5,1.4,1.7]
[null,"837 HOV ",40,1.6025,1.3145121938165154,0.20784262722624466,1.182098647774702,2.0229013522252988,null,1.3,0.7,1.1000000000000001,1.5]

{"sheet":"ctO2","index":109,"n_rows":42,"n_cols":6,"rows_kept":42,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Innkjøring ctO2",null,"Ref. området:                          kalkulert"]
[]
["Dato","Labnummer","735 MB","835 plab","835 MB","837 HOV"]
["2011-10-10",47062477,7.1,6.8,6.9,6.8]
["2011-10-10",47062478,9,8.9,9,8.8000000000000007]
["2011-10-10",47062589,8.1999999999999993,8.1,8.1,8]
["2011-10-10",47062611,5.3,5.2,5.2,5.2]
["2011-10-10",47062649,8.3000000000000007,8.1,8.3000000000000007,8.1]
["2011-10-10",47062666,8.8000000000000007,8.6999999999999993,8.6999999999999993,8.6]
["2011-10-11",47062501,5.4,5.3,5.3,5.4]
["2011-10-11",47062700,6.1,6,6,6.1]
["2011-10-11",47062839,6.7,6.6,6.7,6.6]
["2011-10-12",47063235,5.2,5,5.0999999999999996,5]
["2011-10-12",47063242,7.8,7.6,7.6,7.5]
["2011-10-13",47063363,6.6,6.4,6.2,6.2]
["2011-10-13",47063317,5.9,5.7,5.7,5.8]
["2011-10-13",47063273,5.9,5.9,6,5.9]
["2011-10-13",47063139,5.6,5.6,5.6,5.6]
["2011-10-14",47063445,6.8,6.7,6.7,6.6]
["2011-10-14",47063456,5,4.9000000000000004,4.9000000000000004,4.9000000000000004]
["2011-10-14",47063399,5.8,5.7,5.8,5.7]
["2011-10-14",47063598,8.9,8.6999999999999993,8.8000000000000007,8.6999999999999993]
["2011-10-15",47063791,2.8,2.8,2.8,2.7]
["2011-10-16",47063944,7.4,7.2,7.3,7.2]
["2011-10-16",47063949,7.9,7.8,7.7,7.8]
["2011-10-16",47063979,7.1,6.9,6.9,6.7]
["2011-11-18",47064503,4.9000000000000004,4.8,4.7,4.5999999999999996]
["2011-10-19",47064495,5.2,5.2,5.2,5.3]
["2011-10-19",47064940,8.1,7.9,8,7.8]
["2011-10-20",47065180,7,6.9,6.9,6.8]
["2011-10-20",47065073,3.9,4.2,3.4,4.3]
["2011-10-21",47065432,5.8,6.1,6.1,6.2]
["2011-10-21",47065421,8.6,8.5,8.6,8.5]
["2011-10-26",47066431,8.6,8.4,8.6,8.4]
["2011-10-26",47066582,8,7.9,7.8,7.8]
["2011-10-27",47066702,6.3,6.1,6.2,6.1]
["2011-10-27",47066712,9.1999999999999993,8.9,9.1,8.8000000000000007]
["2011-10-28",47066955,9.3000000000000007,9.1,9.3000000000000007,9.3000000000000007]
["2011-11-01",47067763,6.4,6.2,6.4,6.3]
["2011-11-01",47067808,7.5,7.3,7.4,7.3]
["2011-11-04",47068634,8.6999999999999993,8.5,8.6,8.6]

{"sheet":"tO2 Passing-Bablok725-835plab","index":110,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-14T08:45:34.340"]
[]
[]
[null,"n ",38," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-1.4166666666664902E-2,-0.10000000000000053,0.14999999999999858,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.98333333333333306,0.95833333333333348,1,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","0.05 < p < 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":"(non-linear relationship between X and Y detected)"}]

{"sheet":"tO2 Passing-Bablok725-835MB","index":111,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-14T08:47:48.330"]
[]
[]
[null,"n ",38," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",-9.9999999999999645E-2,-0.10000000000005782,-9.9999999999999645E-2,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":"(constant bias detected)"}]
[null,"Slope ",1,1,1.0000000000000089,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","> 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":""}]

{"sheet":"tO2 Passing-Bablok 725-837HOV","index":112,"n_rows":15,"n_cols":9,"rows_kept":15,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Passing & Bablok method comparison"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,null,"Date ","2011-12-14T08:49:33.740"]
[]
[]
[null,"n ",38," "]
[]
[null,null,"Coefficient","95% CI"]
[null,"Intercept ",4.8571428571429445E-2,-0.19999999999999929,0.30000000000000204,{"$f":"IF(OR(AND(ISNUMBER(D12),D12>0),AND(ISNUMBER(E12),E12<0)),\"(constant bias detected)\",\"\")","v":""}]
[null,"Slope ",0.97142857142857131,0.9375,1,{"$f":"IF(OR(AND(ISNUMBER(D13),D13>1),AND(ISNUMBER(E13),E13<1)),\"(proportional bias detected)\",\"\")","v":""}]
[]
[null,"Cusum test for linearity - p ","0.05 < p < 0.1 ",{"$f":"IF(C15=\"> 0.1 \",\"\", \"(non-linear relationship between X and Y detected)\")","v":"(non-linear relationship between X and Y detected)"}]

{"sheet":"tO2 Bias plot 725-835plab","index":113,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 plab"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-14T08:51:21.330"]
[]
[]
[null,null,null,null,"n ",38," "]
[]
[null,null,null,null,"Bias ",-0.11842105263157894]
[null,null,null,null,"95% CI ",-0.15947127465065364,-7.7370830612504243E-2," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.3632003853368666,-0.43106268457586278,-0.29533808609787043]
[null,null,null,null,"Upper ",0.12635828007370867,5.8495980834712517E-2,0.19422057931270481]

{"sheet":"tO2 Bias plot725-835MB","index":114,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  835 MB"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-14T08:52:47.790"]
[]
[]
[null,null,null,null,"n ",38," "]
[]
[null,null,null,null,"Bias ",-9.210526315789469E-2]
[null,null,null,null,"95% CI ",-0.13491478590109909,-4.9295740414690301E-2," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.34737517145812774,-0.41814586403387116,-0.27660447888238432]
[null,null,null,null,"Upper ",0.16316464514233833,9.2393952566594925E-2,0.23393533771808173]

{"sheet":"tO2 Bias plot 725-837HOV","index":115,"n_rows":16,"n_cols":8,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,"analysed with: Analyse-it + Clinical Laboratory 1.73 "]
[null,"Test "," Bias plots"]
[null,null," "]
[null,null," 735 MB  v  837 HOV"]
[null,"Performed by "," tingstad",null,null,"Date ","2011-12-14T08:54:19.590"]
[]
[]
[null,null,null,null,"n ",38," "]
[]
[null,null,null,null,"Bias ",-0.13421052631578947]
[null,null,null,null,"95% CI ",-0.19102677230474344,-7.7394280326835513E-2," "]
[]
[null,null,null,null,"95% limits of agreement ",null,"95% CI"]
[null,null,null,null,"Lower ",-0.47300145352522804,-0.56692740318628476,-0.37907550386417138]
[null,null,null,null,"Upper ",0.20458040089364904,0.1106544512325924,0.2985063505547057]

{"sheet":"tO2 Descriptives ","index":116,"n_rows":35,"n_cols":13,"rows_kept":35,"cut":false,"hidden":false,"merged":[]}
[]
[null,null,null,null,null,null,null,null,null,null,null,null,"analysed with: Analyse-it + General 1.73 "]
[null,"Test "," Comparative descriptives"]
[null,null," "]
[null,"Variables "," 735 MB, 835 plab, 835 MB, 837 HOV"]
[null,"Performed by "," tingstad",null,null,null,null,null,null,null,"Date ","2011-12-14T08:55:32.820"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null," ","n","Mean","SD","SE","95% CI of Mean",null,null,"Median","IQR","95% CI of Median"]
[null,"735 MB ",38,6.8710526315789462,1.5800971394594772,0.25632560348776906,6.3516876764188233,7.3904175867390691,null,6.9,2.3250000000000002,5.9,7.9]
[null,"835 plab ",38,6.7526315789473692,1.5183093090324964,0.24630229382731769,6.2535757762929762,7.2516873816017622,null,6.75,2.25,6,7.8]
[null,"835 MB ",38,6.7789473684210533,1.5966447750503523,0.25900998444967482,6.2541433413190619,7.3037513955230446,null,6.8,2.2999999999999998,6,7.7]
[null,"837 HOV ",38,6.7368421052631575,1.5129032502660966,0.24542531529151587,6.2395632297347543,7.2341209807915607,null,6.65,2.125,6.1,7.8]